## Dashboard 2018 (hidden)
Sr. No. | TD Section Title | Group | TF-01 | TF-02 | TF-03 | TF-04 | TF-15
1.1.1 | Product or Product Family Name | RA | #REF! | #REF! | #REF! | #REF! | 2019-01-30
1.1.2 | Description of the Device or Device Family | RA | #REF! | #REF! | #REF! | #REF! | 2019-01-30
1.1.3 | Intended use/Purpose | RA | #REF! | #REF! | #REF! | #REF! | 2019-01-30
1.1.4 | Intended Users | RA | #REF! | #REF! | #REF! | #REF! | 2019-01-30
1.1.5 | Product Codes/ Basic UDI-DI | RA | #REF! | #REF! | #REF! | #REF! | 2018-09-30
1.1.6 | Intended Patient Population and Medical Conditions | Med Ops | #REF! | #REF! | #REF! | #REF! | 2019-01-30
1.1.7 | Indications and Contra-Indications | RA | #REF! | #REF! | #REF! | #REF! | 2018-01-30
1.1.8 | Warnings | RA | #REF! | #REF! | #REF! | #REF! | 2019-01-30
1.1.9 | Principles of Operation | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-12-17
1.1.10 | Rationale for the Qualification of the Product as a Device | RA | #REF! | #REF! | #REF! | #REF! | 2018-12-17
1.1.11 | Device Classification and Classification rationale | RA | #REF! | #REF! | #REF! | #REF! | 2018-12-17
1.2.1 | Material Specification | R&D Elements | #REF! | #REF! | #REF! | #REF! | Complete
1.2.1 | Design Specification | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-12-07
1.2.1 | Device Assembly Drawing | R&D Elements | #REF! | #REF! | #REF! | #REF! | Complete
1.2.2 | Novel Features | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-12-07
1.2.3 | Description of the Accessories / Combinations | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-12-07
1.2.4 | Description of Various Configurations / Variants | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-12-07
1.2.5 | Description of Key Functional Elements | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-12-07
1.2.6 | Description of Raw Materials of Key Functional Elements | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-10-15
1.3.1 | Overview of Previous and Similar Generations of the Device | Med Ops | #REF! | #REF! | #REF! | #REF! | 2019-01-30
1.3.2 | Overview of Identified Similar Devices Available on the Markets | Med Ops | #REF! | #REF! | #REF! | #REF! | 2019-01-30
2.1 | Labeling and Packaging Information (Inner / Outer Packaging) | LBL-LBL | #REF! | #REF! | #REF! | #REF! | 2019-01-30
2.2 | Instructions for use | LBL-IFU | #REF! | #REF! | #REF! | #REF! | 2019-01-30
3.1 | Design Stages | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-12-07
3.2 | Manufacturing Process Information | Mfg Eng | #REF! | #REF! | #REF! | #REF! | 2018-12-07
3.3 | Sites including Suppliers and Sub-Contractors | Mfg Eng | #REF! | #REF! | #REF! | #REF! | 2018-12-17
4.1 | General Safety and Performance Requirements Checklist | RA | #REF! | #REF! | #REF! | #REF! | 2019-01-30
4.2 | List of Applied standards | Standards | #REF! | #REF! | #REF! | #REF! | 2018-12-17
5.1 | Benefit-Risk analysis | Risk Mgmt Docs | #REF! | #REF! | #REF! | #REF! | 2019-01-30
5.2 | Risk Management Report / Result | Risk Mgmt Docs | #REF! | #REF! | #REF! | #REF! | 2018-12-07
6.1.1 | Pre-Clinical and Clinical Data Results as well as Design Verification results | Clinical-Reports | #REF! | #REF! | #REF! | #REF! | 2018-12-17
6.1.2 | Biocompatibility | BioComp Testing & Reports | #REF! | #REF! | #REF! | #REF! | 2019-02-15
6.1.3 | Devices Containing CMR or Endocrine-Disrupting Substances | RS Updates | #REF! | #REF! | #REF! | #REF! | 2018-12-17
6.1.4 | Electrical Safety and Electromagnetic Compatibility | R&D Elements | #REF! | #REF! | #REF! | #REF! | N/A
6.1.5 | Software Verification and Validation | R&D Elements | #REF! | #REF! | #REF! | #REF! | N/A
6.1.6 | Device Lifetime | R&D Elements | #REF! | #REF! | #REF! | #REF! | Complete
6.1.7 | Packaging Performance and Transportation | Packaging | #REF! | #REF! | #REF! | #REF! | Complete
6.1.8 | Sterilization | Sterilization Validation/ Reports | #REF! | #REF! | #REF! | #REF! | 2018-12-17
6.1.9 | Summary of Safety and Clinical Performance | SSCP | #REF! | #REF! | #REF! | #REF! | N/A
6.1.10 | Clinical Evaluation Report | Med Ops-CER | #REF! | #REF! | #REF! | #REF! | 2019-01-21
6.1.11 | Post Market Clinical Follow-Up Plan and Report | PMS | #REF! | #REF! | #REF! | #REF! | 2019-01-21
6.2.1 | Devices which Incorporate Medicinal Products | R&D Elements | #REF! | #REF! | #REF! | #REF! | N/A
6.2.2 | Devices which Incorporate Tissues or Cells of Human or Animal Origin or their Derivatives | BioComp Testing & Reports | #REF! | #REF! | #REF! | #REF! | 2018-08-30
6.2.3 | Devices that are Composed of Substances or Combinations of Substances that are Intended to be Introduced into the Human Body and that are Absorbed by or Locally Dispersed in the Human Body) | BioComp Testing & Reports | #REF! | #REF! | #REF! | #REF! | 2018-08-30
6.2.4 | Devices with a Measuring Function | R&D Elements | #REF! | #REF! | #REF! | #REF! | N/A
6.2.5 | Devices Connected with other Devices in order to Operate as Intended | R&D Elements | #REF! | #REF! | #REF! | #REF! | N/A
| Post market documentation | PMS | #REF! | #REF! | #REF! | #REF! | 2018-10-29
| SUBMIT/ACCEPT by RA | RA | #REF! | #REF! | #REF! | #REF! | 2019-02-15
| Latest Date |  | #REF! | #REF! | #REF! | #REF! | 2019-02-15
|  |  | TF-01 (1 SKU) | TF-02 (2 SKU) | TF-03 (3 SKU) | TF-04 (9 SKU) | TF-15 (7 SKU)
| Kick-Off |  | 2018-07-10 | 2018-08-27 | 2018-08-27 | 2018-08-27 | 2018-08-01
| ECR # Create |  | ECR 1394: 8/24/18 | ECR 1394: 8/24/18 | ECR 1394: 8/24/18 | ECR 1394: 8/24/18 | ECR 1403: 8/29/18
| ECR # Submitted to RA |  | 2018-10-25 | 10/25/18 | 10/25/18 | 10/25/18 | TBD
ECR Approved by CRB
Execution Summary
| BioCompatibility | BioComp Testing & Reports | #REF! | #REF! | #REF! | #REF! | 2019-02-15
| Clinical | Reports | #REF! | #REF! | #REF! | #REF! | 2018-12-17
| Labeling | Label/IFU | #REF! | #REF! | #REF! | #REF! | 2019-01-30
| Medical Ops | CER | #REF! | #REF! | #REF! | #REF! | 2019-01-30
| Mfg Engineering | Process Flow | #REF! | #REF! | #REF! | #REF! | 2018-12-17
| Packaging | Package | #REF! | #REF! | #REF! | #REF! | 0
| PMS | Plans/Data | #REF! | #REF! | #REF! | #REF! | 2019-01-21
| Product Stewardship | RS Updates | #REF! | #REF! | #REF! | #REF! | 2018-12-17
| R&D Elements | All | #REF! | #REF! | #REF! | #REF! | 2019-01-30
| Regulatory Affairs | RA | #REF! | #REF! | #REF! | #REF! | 2019-01-30
| Risk Management | Documents | #REF! | #REF! | #REF! | #REF! | 2019-02-15
| SSCP | SSCP | #REF! | #REF! | #REF! | #REF! | 2019-01-21
| Standards | List of Stds | #REF! | #REF! | #REF! | #REF! | 2018-12-17
| Sterilization | Docs/Reports | #REF! | #REF! | #REF! | #REF! | 2019-01-21
Technical Documentation Compiled & Published by RA – LOCK - COMPLETE
| DELIVERY/COMPLETE |  | #REF! | #REF! | #REF! | #REF! | 2019-02-15
9.4: Manually Highlighted based on identified risk.
ESCALATION |  | `
| Risk | TF Impact | DATA Impact | Date Raised | DRI | Mitigation | Escalation Date | Notes
1 | Restricted Substance Data Needed from Supplier | TF-01, 02., 03, 04 | RS Data needed for Assessment/ Testing identification | 2018-08-15 | Rupak Chatterjee | None | 2018-08-15 | 9.4: Escalation to SRM, SQE, Procurement.
2 | Manufacturing Process (MVP) Requirements are unclear | Work is pending until clarity is provided. | Work is pending until clarity is provided. | 2018-08-15 | Eric Hulse |  | 2018-08-15
3 | Bio-Compatibility Resources are constrained | Work cannot complete | Work cannot complete | 2018-08-15 | Ed Reverdy |  | 2018-08-15
4 | R&D Design Verification/ Validation Requirements Unclear | ALL | May need to perform testing. | 2018-08-01 | S. Parihar |  | 2018-08-15 | 8.15: Escalated to K. Doll for Resolution/ input/ direction
5 | DPM requirements are unknown; Need NB number | Delay DPM on device | Delay DPM on device | 2018-08-15 | Eric Hulse |  | 2018-08-15
6 | Bio-Comp Testing will require 5-6 Months | TF-15 | Down Stream reporting | 2018-08-01 | B. Przygoda | None | 2018-08-15 | Samples are in-hand; testing will begin 8/29

## 2018 SUMRY (hidden)
SUMMARY TAB
Sr. No. | Document | Responsible Function | EH40 |  | Comments / Remarks | Duration (day) | Start Date | Completion date | Deliverable | Completion Date TF-01 | Completion Date TF-02 | Completion Date TF-03 | Completion Date TF-04 | Completion Date TF-05
|  |  | From (Current Technical File TF N01) | To (As per MDR 2017/745)
1 | DEVICE DESCRIPTION AND SPECIFICATION, INCLUDING VARIANTS AND ACCESSORIES | [7 empty] | Deliverable
1.1 | Description of Device or Device Family | [7 empty] | Section Header | Section Header | Section Header | Section Header | Section Header | Section Header
1.1.1 | Product or Product Family Name | RA | Section 1: Product Information | Purse String Clamp | Guidance/description : product or trade name as per product label and Product Family Name as appropriate. Source: Technical File (TF - N01) | 1 |  | 2018-09-30 | RA | #REF! | #REF! | #REF! | #REF! | 2019-01-30
1.1.2 | Description of the Device or Device Family | RA | Section 1.2 : General Product Description | Content in Technical Documentation (TD) to be finalized | Guidance/description : General description from the latest verified source of information and include a full description to ensure reviewer can understand the device as fully as possible. Suggested Source : Technical File (TF - N01), Design Specification (DR000431), CER (SCN048339), IFU etc. | 1 |  | 2018-09-30 | RA | #REF! | #REF! | #REF! | #REF! | 2019-01-30
1.1.3 | Intended use/Purpose | RA, Medical Affairs | N/A | New Section: Content in TD to be finalized | Guidance/description : “Intended Purpose” means the use for which a device is intended according to the data supplied by the manufacturer on the label, IFU or in promotional or sales materials or statements and as specified by the manufacturer in the CER. Suggested Source : CER (SCN048339), IFU etc. | 1 |  | 2018-12-14 | RA | #REF! | #REF! | #REF! | #REF! | 2019-01-30
1.1.4 | Intended Users | RA | N/A | New Section: Content in TD to be finalized | Guidance/description : Intended User information can be used from IFU and CER SCN048339, when CER will be updated as per MDR requirements | 1 |  | 2018-12-14 | RA | #REF! | #REF! | #REF! | #REF! | 2019-01-30
1.1.5 | Product Codes/ Basic UDI-DI | RA | N/A | New Section: Content in TD to be finalized | Guidance/description : Device UDI number to be mentioned in this section. Suggested Source : GSI database GTIN (EH40) : 10705031135208 | 1 |  | 2018-08-10 | RA | #REF! | #REF! | #REF! | #REF! | 2018-09-30
1.1.6 | Intended Patient Population and Medical Conditions | RA, Med Affairs, QE | N/A | New Section: Content in TD to be finalized | Guidance/description : Intended patient population, medical conditions to be diagnosed, treated and/or monitored and patient selection criteria can be located in CER SCN048339, when CER will be updated as per MDR requirements | 2 |  | 2018-12-14 | Med Ops | #REF! | #REF! | #REF! | #REF! | 2019-01-30
1.1.7 | Indications and Contra-Indications | RA | N/A | New Section: Content in TD to be finalized | Guidance/description : Indications for use and other considerations such as contra-indications can be located in the Instructions for Use (IFU). | 1 |  | 2018-12-14 | RA | #REF! | #REF! | #REF! | #REF! | 2018-01-30
1.1.8 | Warnings | RA | N/A | New Section: Content in TD to be finalized | Guidance/description : Warnings can be located in the Instructions for Use (IFU). This should also include considerations based on devices construction, uses, procedures and associated risk. | 1 |  | 2018-12-14 | RA | #REF! | #REF! | #REF! | #REF! | 2019-01-30
1.1.9 | Principles of Operation | RA, R&D | N/A | New Section: Content in TD to be finalized | Guidance/description : Principles of operation of the device and its mode of action shall be provided with scientifically demonstration if necessary. The mode of action should clearly identify how the device works. Suggested Source : Design Specification (DR000431), CER (SCN048339) , IFU etc. | 1 |  | 2018-12-14 | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-12-17
1.1.10 | Rationale for the Qualification of the Product as a Device | RA | N/A | New Section: Content in TD to be finalized | Guidance/description : A rationale to be provided for the qualification of the product as a device that can be located under applicable definition from MDR 2017/745 Article 2 (Definitions). | 1 |  | 2018-12-14 | RA | #REF! | #REF! | #REF! | #REF! | 2018-12-17
1.1.11 | Device Classification and Classification rationale | RA | N/A | New Section: Content in TD to be finalized | Guidance/description : Review Annex VIII of MDR 2017/745 to determine the classification and applicable rule(s). Address other applicable classification rules: Ensure to document why considered rules were eliminated for the product. Provide information as to whether the device is a Well-Established Technology (WET) device or provide justification if not. Ensure consistency with CER (SCN048339), when CER will be updated as per MDR requirements | 1 |  | 2018-12-14 | RA | #REF! | #REF! | #REF! | #REF! | 2018-12-17
1.2 | Technical Specification | [7 empty] | Section Header | Section Header | Section Header | Section Header | Section Header | Section Header
1.2.1 | Product Specification: | [7 empty] | Section Header | Section Header | Section Header | Section Header | Section Header | Section Header
| Material Specification | R&D, Med Operations, Med Affairs, Clinical | MS8087C | MS8087C Rev. XXX | No gaps identified other than requirements for performing new inspections if UDI marking is to be introduced for Purse String Clamp | 2 | 2018-08-01 | 2018-08-03 | R&D Elements | #REF! | #REF! | #REF! | #REF! | Complete
| Design Specification | R&D, Med Operations, Med Affairs, Clinical | DR000431 | DR000431 | No gaps identified after review | 0 |  | Complete | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-12-07
| Device Assembly Drawing | R&D, Med Operations, Med Affairs, Clinical | D44914 / D03753 | D44914 / D03753 | No gaps identified after review | 0 |  | Complete | R&D Elements | #REF! | #REF! | #REF! | #REF! | Complete
1.2.2 | Novel Features | R&D, Med Operations, Med Affairs, Clinical | N/A | New Section: Content in TD to be finalized | Guidance/description : Provide an explanation of any novel features e.g. innovative features, new to the market, based on claims and design input/output. If there are no novel features, state this and provide justification. Suggested Source : CER (SCN048339) and design intent analysis | 10 | 2018-08-06 | 2018-08-16 | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-12-07
1.2.3 | Description of the Accessories / Combinations | R&D | N/A | New Section: Content in TD to be finalized | Guidance/description : Provide a description of the accessories that are intended to be used in combination with the device (interconnectivity). State whether the accessories are provided with the device or if they are not included. Suggested Source : Information taken from CER (SCN048339): section 4.2, The PURSE STRING CLAMP requires the use of needle and suture per user preference. | 1 | 2018-08-01 | 2018-08-03 | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-12-07
1.2.4 | Description of Various Configurations / Variants | R&D | N/A | New Section: Content in TD to be finalized | Guidance/description : A description of the accessories that are intended to be used in combination with the device (interconnectivity). State whether the accessories are provided with the device or if they are not included. This section should also indicate how the device is sold and used. | 1 | 2018-08-20 | 2018-08-21 | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-12-07
1.2.5 | Description of Key Functional Elements | R&D | N/A | New Section: Content in TD to be finalized | Guidance/description : A description of the key functional elements, e.g. the device parts/components the formulation, composition, functionality and, where relevant, qualitative and quantitative composition (i.e. what parts and what material is each part made of). Where appropriate, this shall include labelled pictorial representations (e.g. diagrams, photographs, and drawings), clearly indicating key components, including sufficient explanation to understand the drawings and diagrams. | 15 | 2018-08-20 | 2018-09-07 | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-12-07
1.2.6 | Description of Raw Materials of Key Functional Elements | R&D, Biocompatibility | N/A | New Section: To be obtained from DBAM | Guidance/description : A description of the raw materials incorporated into key functional elements and those making either direct contact with the human body or indirect contact with the body, e.g. during extracorporeal circulation of body fluids. | 2 | 2018-09-10 | 2018-09-12 | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-10-15
1.3 | Overview of Previous and Similar Generations of the Device | [7 empty] | Section Header | Section Header | Section Header | Section Header | Section Header | Section Header
1.3.1 | Overview of Previous and Similar Generations of the Device | RA, R&D, Medical Operations, Clinical | N/A | New Section: Content in TD to be finalized | Guidance/description : Provide overview of the previous generation(s) of the device by the manufacturer, where such devices exist. Suggested Source : CER (SCN048339) can be referred, when CER will be updated as per MDR requirements | 20 | 2018-09-01 | 2018-09-30 | Med Ops | #REF! | #REF! | #REF! | #REF! | 2019-01-30
1.3.2 | Overview of Identified Similar Devices Available on the Markets | RA, Medical Operations, Clinical | N/A | New Section: Content in TD to be finalized | Guidance/description : Provide an overview of identified similar devices available in the Union or international markets, where such devices exist. Suggested Source : CER (SCN048339) can be referred, when CER will be updated as per MDR requirements | 20 | 2018-09-01 | 2018-09-30 | Med Ops | #REF! | #REF! | #REF! | #REF! | 2019-01-30
2 | INFORMATION TO BE SUPPLIED BY THE MANUFACTURER | [7 empty] | Section Header | Section Header | Section Header | Section Header | Section Header | Section Header
2.1 | Labeling and Packaging Information (Inner / Outer Packaging) | Labeling, Packaging, RA | Package Assembly Drawings P000432 Generic Carton Drawings P000520 (P000422P01 / A000805P00) | Package Assembly Drawings P000432 Generic Carton Drawings P000520 (P000422P01 / A000805P00 (Working stage in EPI)) | Based on New MDR requirements, package assembly and Generic Carton drawings may require further modification due to UDI implementation. Require SME input. |  |  | 2018-09-30 | LBL-LBL | #REF! | #REF! | #REF! | #REF! | 2019-01-30
2.2 | Instructions for use | Labeling, RA | Current Production IFU SCN053725 (RMC8340455) | New document as per CLDP process | IFU to be updated by including information based on new requirements for MDR such as : 1) Patient Target Group 2) Clinical Benefits 3) Safety and Clinical Performance (SSCP) (if applicable) 4) Qualitative and quantitative information on implant materials 5) Warning against reporting of any serious incident etc. |  |  | 2018-09-30 | LBL-IFU | #REF! | #REF! | #REF! | #REF! | 2019-01-30
3 | DESIGN AND MANUFACTURING INFORMATION | [7 empty] | Section Header | Section Header | Section Header | Section Header | Section Header | Section Header
3.1 | Design Stages | R&D, QE, Manufacturing Engineering | N/A | New Section: Content in TD to be finalized | Reference to various design factbooks (DHFs) as available for Purse String Clamp is required to be filled in this section as well as applicable EES procedure need to be depicted in this section |  |  | 2018-09-30 | R&D Elements | #REF! | #REF! | #REF! | #REF! | 2018-12-07
3.2 | Manufacturing Process Information | R&D, QE, Manufacturing Engineering | Detail Manufacturing Information including manufacturing flow chart and process specification SCN049769 | Detail Manufacturing Information including manufacturing flow chart and process specification SCN049769 | Other than Manufacturing process validations(Master Validation Plan), all required documents are available which comply with MDR requirements |  |  | 2018-10-30 | Mfg Eng | #REF! | #REF! | #REF! | #REF! | 2018-12-07
3.3 | Sites including Suppliers and Sub-Contractors | R&D, RA, QE, Manufacturing Engineering, Supplier Quality | N/A | New Section: Content in TD to be finalized | Information of all suppliers related to key aspects of a product such as : design, raw material manufacturing, packaging, sterilization etc. are required in this section. Also need to provide Authorized Representative in this section. |  |  | 2018-09-30 | Mfg Eng | #REF! | #REF! | #REF! | #REF! | 2018-12-17
4 | GENERAL SAFETY AND PERFORMANCE REQUIREMENTS | [7 empty] | Section Header | Section Header | Section Header | Section Header | Section Header | Section Header
4.1 | General Safety and Performance Requirements Checklist | RA | N/A | 100643298 General Safety and Performance Requirements Checklist SCNXXXXXX | Guidance/description : The device must meet the General Safety and Performance Requirements set out in Annex I of the Medical Device Regulation (EU) 2017/745, which apply to it, taking account of the intended purpose of the device concerned. The compliance of the device to the relevant requirements shall be documented in the referenced checklist(s). A justification is provided within the checklist for the requirements that do not apply. |  |  |  | RA | #REF! | #REF! | #REF! | #REF! | 2019-01-30
4.2 | List of Applied standards | Standards | SCN053711 | SCN053711 rev. XXX | Document to be revised considering new list of applicable Harmonized Standards as applicable in EES after MDR implementation |  |  | 2018-09-30 | Standards | #REF! | #REF! | #REF! | #REF! | 2018-12-17
5 | BENEFIT-RISK ANALYSIS AND RISK MANAGEMENT | [7 empty] | Section Header | Section Header | Section Header | Section Header | Section Header | Section Header
5.1 | Benefit-Risk analysis | QE, Risk Mgmt. | N/A | New Section: Content in TD to be finalized | Information related to risk benefit analysis can be obtained from the following sources Suggested Source : CER (SCN048339), Risk management Report (RMD000635) |  |  |  | Risk Mgmt Docs | #REF! | #REF! | #REF! | #REF! | 2019-01-30
5.2 | Risk Management Report / Result | QE, Risk Mgmt. | RMR RMD000635 Process FMEA (device) RMD001344 Process FMEA (packaging) SCN047340 Design FMEA (device) RMD001341 Design FMEA (packaging) SCN047339 Product Use Risk Analysis (PURA4R) RMD001384 Application FMEA RMD000621 | RMR RMD000635 Rev. XXX Process FMEA (device) RMD001344 Rev. XXX Process FMEA (packaging) SCN047340 Design FMEA (device) RMD001341 Rev XXX Design FMEA (packaging) SCN047339 Product Use Risk Analysis (PURA4R) RMD001384 Rev XXX Application FMEA RMD000621 Rev XXX | Modify Risk management documents as per new hazards identified based on 103433773 MDD to MDR Compliance Procedure |  |  | 2018-09-30 | Risk Mgmt Docs | #REF! | #REF! | #REF! | #REF! | 2018-12-07
6 | PRODUCT VERIFICATION AND VALIDATION | [7 empty] | Section Header | Section Header | Section Header | Section Header | Section Header | Section Header
6.1 | Pre-Clinical and Clinical Data Result | [7 empty] | Section Header | Section Header | Section Header | Section Header | Section Header | Section Header
6.1.1 | Pre-Clinical and Clinical Data Results as well as Design Verification results | Pre-Clinical, Clinical, R&D | CER - SCN048339 | New Section: Content in TD to be finalized | Guidance/description : Provide the detailed information regarding test design, test protocol, methods of data analysis, in addition to data summaries and test conclusions regarding each of the pre-clinical evaluation activities. May need to provide Justification memo in lack of design verification testing for legacy devices in market for more that 15 years. Suggested Source: CER (SCN048339), Design Verification Matrix (DR000432), Design Factbook (FB000331) etc. |  |  | 10/15/18 (11/30/18 if testing required) | Clinical-Reports | #REF! | #REF! | #REF! | #REF! | 2018-12-17
6.1.2 | Biocompatibility | Biocompatibility | SCN046114 | SCN046114 Rev. XXX | Existing DBAM document to be updated to capture the new R-Spec numbers pulled based on MDR requirements Suggested Source: DBAM: (SCN046114) |  |  | 2018-11-30 | BioComp Testing & Reports | #REF! | #REF! | #REF! | #REF! | 2019-02-15
6.1.3 | Devices Containing CMR or Endocrine-Disrupting Substances | Product Stewardship | N/A | New Section: Content in TD to be finalized | Guidance/description : The device component materials shall be reviewed. This review shall be based upon supplier material disclosures, laboratory testing and screening, and technical assessments. | 5 | 2018-09-17 | 2018-09-21 | RS Updates | #REF! | #REF! | #REF! | #REF! | 2018-12-17
6.1.4 | Electrical Safety and Electromagnetic Compatibility | R&D | N/A | New Section: Not Applicable | Since Purse String Clamp does not have any medical electrical equipment and medical electrical systems, this section will not be applicable. Hence no gaps identified. | 0 |  |  | R&D Elements | #REF! | #REF! | #REF! | #REF! | N/A
6.1.5 | Software Verification and Validation | R&D | N/A | New Section: Not Applicable | Since Purse String Clamp does not have any software related systems, this section will not be applicable. Hence no gaps identified. | 0 |  |  | R&D Elements | #REF! | #REF! | #REF! | #REF! | N/A
6.1.6 | Device Lifetime | LM, R&D | Stability Strategy (device) DOC022849 | Stability Strategy (device) DOC022849 Rev. XXX | MDR requires device shall mention its intended lifetime within existing applicable documentation. | 5 |  | 2018-09-15 | R&D Elements | #REF! | #REF! | #REF! | #REF! | Complete
6.1.7 | Packaging Performance and Transportation | Packaging, R&D | Package Validation (Transit test) SCN046113 | Package Validation (Transit test) SCN046113 | No gaps identified |  |  | 2018-08-30 | Packaging | #REF! | #REF! | #REF! | #REF! | Complete
6.1.8 | Sterilization | Sterilization | Sterilization strategy SCN046313 | Sterilization strategy SCN046313 Rev. XXX | Sterilization group to provide feedback on any modification if necessary for the Purse String Clamp existing sterilization strategy |  |  | 2018-08-30 | Sterilization Validation/ Reports | #REF! | #REF! | #REF! | #REF! | 2018-12-17
6.1.9 | Summary of Safety and Clinical Performance | Med Affairs, Med Ops, Clinical | N/A | New Section: Not Applicable | A Summary of Safety and Clinical Performance (SSCP) is required for implantable devices and class III devices & not for class I devices | 0 |  |  | SSCP | #REF! | #REF! | #REF! | #REF! | N/A
6.1.10 | Clinical Evaluation Report | Med Affairs, Med Ops, Clinical | CER - SCN048339 | CER - SCN048339 Rev. XXX | CER to be updated based on MDR requirements | 15 | 2018-10-29 | 2018-12-13 | Med Ops-CER | #REF! | #REF! | #REF! | #REF! | 2019-01-21
6.1.11 | Post Market Clinical Follow-Up Plan and Report | PMS, Med Affairs, Clinical | PMS Plan - RMD001028 PMS Reports data for CER -RMD001136 | PMS Plan - RMD001028 PMS Report - RMD001136 | Guidance/description : The PMCF is a continuous process that updates the clinical evaluation and should be addressed in the Post-Market Surveillance Plan. Suggested Source: PMS Plan & PMS Report; CER (SCN048339), when CER will be updated based on MDR requirements |  |  | 2018-10-28 | PMS | #REF! | #REF! | #REF! | #REF! | 2019-01-21
6.2 | Additional Information Required in Specific Cases | [7 empty] | Section Header | Section Header | Section Header | Section Header | Section Header | Section Header
6.2.1 | Devices which Incorporate Medicinal Products | R&D | N/A | New Section: Applicability in TD to be finalized | Guidance/description : This section refers to the subject device incorporates, as an integral part, a substance which, if used separately, within the meaning of point 2 of Article 1 of Directive 2001/83/EC, including a medicinal product derived from human blood or human plasma, is considered to be a medicinal product. | 0 |  |  | R&D Elements | #REF! | #REF! | #REF! | #REF! | N/A
6.2.2 | Devices which Incorporate Tissues or Cells of Human or Animal Origin or their Derivatives | Biocompatibility | N/A | New Section: Applicability in TD to be finalized | Guidance/description : This section refers to the subject device if manufactured utilizing tissues or cells of human or animal origin, or their derivatives covered by MDR, point (f) and (g) of Article 1(6), and / or Article 1(10). | 0 |  |  | BioComp Testing & Reports | #REF! | #REF! | #REF! | #REF! | 2018-08-30
6.2.3 | Devices that are Composed of Substances or Combinations of Substances that are Intended to be Introduced into the Human Body and that are Absorbed by or Locally Dispersed in the Human Body) | Biocompatibility | N/A | New Section: Applicability in TD to be finalized | Guidance/description : This section refers to the subject device if composed of substances or combinations of substances that are intended to be introduced into the human body and are absorbed by or locally dispersed in the human body. | 0 |  |  | BioComp Testing & Reports | #REF! | #REF! | #REF! | #REF! | 2018-08-30
6.2.4 | Devices with a Measuring Function | R&D | N/A | New Section: Applicability in TD to be finalized | Guidance/description : This section refers to a description of the methods used to ensure the accuracy as given in the specifications. | 0 |  |  | R&D Elements | #REF! | #REF! | #REF! | #REF! | N/A
6.2.5 | Devices Connected with other Devices in order to Operate as Intended | R&D | N/A | New Section: Content in TD to be finalized | Guidance/description : The documentation referenced shall demonstrate that the subject device is connected to other device(s) during operation as intended, and conforms to the General Safety and Performance Requirements. | 0 |  |  | R&D Elements | #REF! | #REF! | #REF! | #REF! | N/A
7 | POST MARKET SURVEILLANCE | [7 empty] | Section Header | Section Header | Section Header | Section Header | Section Header | Section Header
| Post market documentation | PMS | PMS Plan - RMD001036 PMS Reports data - RMD000830 | PMS Plan - RMD001036 Rev. XXX PMS Reports data - RMD000830 Rev. XXX | PMS plan and Report to be finalized based on MDR requirements |  |  | 2018-10-28 | PMS | #REF! | #REF! | #REF! | #REF! | 2018-10-29
8 | EU DECLARATION OF CONFORMITY | [7 empty] | Section Header | Section Header | Section Header | Section Header | Section Header | Section Header
| SUBMIT/ACCEPT by RA | RA | N/A | 100642396 MDR EU Declaration of Conformity Form (SCNXXXXXX - Purse String Clamp) | Document to be created based on MDR requirements |  |  | 2018-12-18 | RA | #REF! | #REF! | #REF! | #REF! | 2019-02-15
9 | CERTIFICATES | [7 empty] | Section Header | Section Header | Section Header | Section Header | Section Header | Section Header
9.1 | EU Certificate | RA | N/A | XXXX | Certificates to be obtained for MDR compliance |  |  |  | OOS | OOS | OOS | OOS | OOS | OOS
9.2 | EN ISO 13485 Certificate | RA, QA | N/A | XXXX |  |  |  |  | OOS | OOS | OOS | OOS | OOS | OOS
| Latest Date | [8 empty] | #REF! | #REF! | #REF! | #REF! | 2019-02-15

## LT Clips (hidden)
LIGACLIP EXTRA Ligating Clip Cartridge Titanium Clips (Class IIB)
Sr. No. | Document | Responsible Function |  |  | Comments / Remarks | Deliverables | Dependency (Source of Information) | Status | Remarks | Duration (day) | Start Date | Dates without Dependency | Estimated Completion date | Associated Dependencies | Deliverable
|  |  | From (Current Technical File TF N000006) | To (As per MDR 2017/745)
1 | DEVICE DESCRIPTION AND SPECIFICATION, INCLUDING VARIANTS AND ACCESSORIES | [13 empty] | Deliverable
1.1 | Description of Device or Device Family | [13 empty] | Section Header
1.1.1 | Product or Product Family Name | RA |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs to update current IFU as per EU MDR |  |  | 2018-09-30 | 2019-01-30 |  | RA
1.1.2 | Description of the Device or Device Family | RA |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs |  |  | 2018-09-30 | 2019-01-30 |  | RA
1.1.3 | Intended use/Purpose | RA, Medical Affairs |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs |  |  | 2019-01-30 | 2019-01-30 | Need CER alignment | RA
1.1.4 | Intended Users | RA |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Pending Information | Current IFU donot list intended users information in Warnings section Labelling team to finalize information with other functional SMEs |  |  | 2019-01-30 | 2019-01-30 | Need CER alignment | RA
1.1.5 | Product Codes/ Basic UDI-DI | RA |  |  | Suggested Source : GSI database GTIN |  | N/A | Pending Information | Need information for UDI |  |  | 2018-09-30 | 2018-09-30 |  | RA
1.1.6 | Intended Patient Population and Medical Conditions | RA, Med Affairs, QE |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Pending Information | Current IFU do not list intended Patient Population and Medical Conditions information Need to update |  |  | 2019-01-30 | 2019-01-30 | Need CER alignment | Med Ops
1.1.7 | Indications and Contra-Indications | RA |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs |  |  | 2018-01-30 | 2018-01-30 | Need CER alignment | RA
1.1.8 | Warnings | RA |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs |  |  | 2019-01-30 | 2019-01-30 | Need CER alignment | RA
1.1.9 | Principles of Operation | RA, R&D |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-10-15 | 2018-12-17 |  | R&D Elements
1.1.10 | Rationale for the Qualification of the Product as a Device | RA |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-10-30 | 2018-12-17 |  | RA
1.1.11 | Device Classification and Classification rationale | RA |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-10-30 | 2018-12-17 |  | RA
1.2 | Technical Specification | [13 empty] | Section Header
1.2.1 | Product Specification: | [13 empty] | Section Header
| Material Specification | R&D, Med Operations, Med Affairs, Clinical |  |  |  |  | N/A | TD Deliverables Released |  |  |  |  | Complete |  | R&D Elements
| Design Specification | R&D, Med Operations, Med Affairs, Clinical | [5 empty] | Pending Review with SME | PSX - DRXXXXXX to be released in EPIcenter after review |  |  | 2018-10-15 | 2018-12-07 |  | R&D Elements
| Device Assembly Drawing | R&D, Med Operations, Med Affairs, Clinical |  |  |  |  | N/A | TD Deliverables Released |  |  |  |  | Complete |  | R&D Elements
1.2.2 | Novel Features | R&D, Med Operations, Med Affairs, Clinical |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-11-30 | 2018-12-07 |  | R&D Elements
1.2.3 | Description of the Accessories / Combinations | R&D |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Pending Review with SME |  |  |  | 2015-09-30 | 2018-12-07 |  | R&D Elements
1.2.4 | Description of Various Configurations / Variants | R&D |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-09-30 | 2018-12-07 |  | R&D Elements
1.2.5 | Description of Key Functional Elements | R&D | N/A | New Section: Content in TD to be finalized with KFE Memo (SCNXXXXX) | Guidance/description : A description of the key functional elements, e.g. the device parts/components the formulation, composition, functionality and, where relevant, qualitative and quantitative composition (i.e. what parts and what material is each part made of). Where appropriate, this shall include labelled pictorial representations (e.g. diagrams, photographs, and drawings), clearly indicating key components, including sufficient explanation to understand the drawings and diagrams. |  | DBAM - SCN018493 Rev. XXX Design Specification - PMR35 - FB000551 p. 33 PMW35 - FB000551 p. 33 PRR35 - FB000501 p. 22 PRW35 - FB000501 p. 22 PXR35 - DS0057 PXW35 - DS0057 PSX - DRXXXXXX Sterilization Strategy - SCN005386, SCN008851 | Pending Review with SME | ●DBAM to be updated based on latest test requirements |  |  | 2018-10-15 | 2018-12-07 | Need Device testing | R&D Elements
1.2.6 | Description of Raw Materials of Key Functional Elements | R&D, Biocompatibility |  |  |  |  | DBAM - SCN018493 Rev. XXX | Pending Review with SME | ●DBAM to be updated based on latest test requirements ●KFE to be released after finalizaton of other sections within the document |  |  | 2018-10-15 | 2018-10-15 | Need Device testing | R&D Elements
1.3 | Overview of Previous and Similar Generations of the Device | [13 empty] | Section Header
1.3.1 | Overview of Previous and Similar Generations of the Device | RA, R&D, Medical Operations, Clinical |  |  |  |  | CER - SCN019399 | Pending Information | Need updates based on latest CER after MDR related updates to CER |  |  | 2018-11-30 | 2019-01-30 |  | Med Ops
1.3.2 | Overview of Identified Similar Devices Available on the Markets | RA, Medical Operations, Clinical |  |  |  |  | CER - SCN019399 | Pending Information | Need updates based on latest CER after MDR related updates to CER |  |  | 2018-11-30 | 2019-01-30 |  | Med Ops
2 | INFORMATION TO BE SUPPLIED BY THE MANUFACTURER | [13 empty] | Section Header
2.1 | Labeling and Packaging Information (Inner / Outer Packaging) | Labeling, Packaging, RA |  |  |  |  | Explained in remarks column | Pending Information | May need updates based on latest MDR related updates to label information |  |  | 2018-11-01 | 2019-01-30 |  | LBL-LBL
2.2 | Instructions for use | Labeling, RA |  |  |  |  | Explained in remarks column | Pending Information | May need updates based on latest MDR related updates to include information for referred TD sections : 1.1.4 1.1.6 and 1.2.3 |  |  | 2018-11-01 | 2019-01-30 |  | LBL-IFU
3 | DESIGN AND MANUFACTURING INFORMATION | [13 empty] | Section Header
3.1 | Design Stages | R&D, QE, Manufacturing Engineering |  |  |  |  | Explained in remarks column | Working | Design Specification - PSX will be released after review Design Verification Matrix - PMR and PMW is being created and will be released after review ● Updated Justification Memo to be released |  |  | 2018-10-30 | 2018-12-07 |  | R&D Elements
3.2 | Manufacturing Process Information | R&D, QE, Manufacturing Engineering |  |  |  |  | Explained in remarks column | Pending Review with SME | Manufacturing Memo for Skin Staplers and Extractor - SCNXXXXXXX to be released after review with SME |  |  | 2018-11-20 | 2018-12-07 |  | Mfg Eng
3.3 | Sites including Suppliers and Sub-Contractors | R&D, RA, QE, Manufacturing Engineering, Supplier Quality |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-09-30 | 2018-12-17 |  | Mfg Eng
4 | GENERAL SAFETY AND PERFORMANCE REQUIREMENTS | [13 empty] | Section Header
4.1 | General Safety and Performance Requirements Checklist | RA |  |  |  |  | Explained in remarks column | Pending Review with SME | Depends on all deliverables |  |  | 2018-11-30 | 2019-01-30 |  | RA
4.2 | List of Applied standards | Standards |  |  |  |  | Explained in remarks column | Pending Review with SME | All Functional groups to conform to latest standard requirements |  |  | 2018-10-30 | 2018-12-17 |  | Standards
5 | BENEFIT-RISK ANALYSIS AND RISK MANAGEMENT | [13 empty] | Section Header
5.1 | Benefit-Risk analysis | QE, Risk Mgmt. |  |  |  |  | N/A | Pending Review with SME | May need updates based on latest CER after MDR related updates to CER |  |  | 2018-11-15 | 2019-01-30 |  | Risk Mgmt Docs
5.2 | Risk Management Report / Result | QE, Risk Mgmt. |  |  |  |  | Explained in remarks column | Working | Hazard Analysis Memo - SCNXXXXX Application FMEA - RMD001200, RMD001189 Design FMEA - DF0120, RMD001203, RMD001188 Documents mentioned above to be released after SME review |  |  | 2018-11-15 | 2018-12-07 |  | Risk Mgmt Docs
6 | PRODUCT VERIFICATION AND VALIDATION | [13 empty] | Section Header
6.1 | Pre-Clinical and Clinical Data Result | [13 empty] | Section Header
6.1.1 | Pre-Clinical and Clinical Data Results as well as Design Verification results | Pre-Clinical, Clinical, R&D |  |  |  |  | Explained in remarks column | Working | Design Verifiation Matrix - PMR and PMW to be reviewed by SME along with associated Preclinical information Above documents to be released after review |  |  | 2018-11-30 | 2018-12-17 |  | Clinical-Reports
6.1.2 | Biocompatibility | Biocompatibility |  |  |  |  | N/A | Working | DBAM to be updated based on latest Biocompatibility Testing standards |  |  | 2019-02-15 | 2019-02-15 | Need Device testing | BioComp Testing & Reports
6.1.3 | Devices Containing CMR or Endocrine-Disrupting Substances | Product stewardship |  |  |  |  | N/A | Working |  |  |  | 2018-10-15 | 2018-12-17 |  | RS Updates
6.1.4 | Electrical Safety and Electromagnetic Compatibility | R&D |  |  |  |  | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
6.1.5 | Software Verification and Validation | R&D |  |  |  |  | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
6.1.6 | Device Lifetime | LM, R&D |  |  |  |  | N/A | TD Deliverables Released |  |  |  | 2018-09-15 | Complete |  | R&D Elements
6.1.7 | Packaging Performance and Transportation | Packaging, R&D |  |  |  |  | N/A | TD Deliverables Released |  |  |  | 2018-10-31 | Complete |  | Packaging
6.1.8 | Sterilization | Sterilization |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-10-30 | 2018-12-17 |  | Sterilization Validation/ Reports
6.1.9 | Summary of Safety and Clinical Performance | Med Affairs, Med Ops, Clinical |  |  |  |  | N/A | N/A |  |  |  | 2018-10-30 | N/A |  | SSCP
6.1.10 | Clinical Evaluation Report | Med Affairs, Med Ops, Clinical |  |  |  |  | Explained in remarks column | Working | ●R&D to Provide information to Principle of Operation section of CER ●Update relavant information such as intended patient population medical condition; literature study report etc. as applicable ● CER should be consistent with the information as provided within IFU and Technical documentation sections where applicable Example: May need updates based on latest MDR related updates to include information for referred TD sections: 1.3.1 and 1.3.2 |  |  | 2019-01-21 | 2019-01-21 | Need Biocompatibility input | CER
6.1.11 | Post Market Clinical Follow-Up Plan and Report | PMS, Med Affairs, Clinical |  |  |  |  | PMS Plan - RMD001036 PMS Reports data for CER - RMD000830 | Pending Review with SME | May need updates based on latest CER after MDR related updates to CER |  |  | 2018-10-29 | 2019-01-21 |  | PMS
6.2 | Additional Information Required in Specific Cases | [13 empty] | Section Header
6.2.1 | Devices which Incorporate Medicinal Products | R&D | N/A | New Section: Not Applicable | Guidance/description : This section refers to the subject device incorporates, as an integral part, a substance which, if used separately, within the meaning of point 2 of Article 1 of Directive 2001/83/EC, including a medicinal product derived from human blood or human plasma, is considered to be a medicinal product. |  | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
6.2.2 | Devices which Incorporate Tissues or Cells of Human or Animal Origin or their Derivatives | Biocompatibility | N/A | New Section: Not Applicable | Guidance/description : This section refers to the subject device if manufactured utilizing tissues or cells of human or animal origin, or their derivatives covered by MDR, point (f) and (g) of Article 1(6), and / or Article 1(10). |  | N/A | Pending Review with SME |  |  |  | 2018-08-30 | 2018-08-30 |  | BioComp Testing & Reports
6.2.3 | Devices that are Composed of Substances or Combinations of Substances that are Intended to be Introduced into the Human Body and that are Absorbed by or Locally Dispersed in the Human Body) | Biocompatibility | N/A | New Section: Not Applicable | Guidance/description : This section refers to the subject device if composed of substances or combinations of substances that are intended to be introduced into the human body and are absorbed by or locally dispersed in the human body. |  | N/A | Pending Review with SME |  |  |  | 2018-08-30 | 2018-08-30 |  | BioComp Testing & Reports
6.2.4 | Devices with a Measuring Function | R&D | N/A | New Section: Not Applicable | Guidance/description : This section refers to a description of the methods used to ensure the accuracy as given in the specifications. |  | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
6.2.5 | Devices Connected with other Devices in order to Operate as Intended | R&D | N/A | New Section: Content in TD to be finalized | Guidance/description : The documentation referenced shall demonstrate that the subject device is connected to other device(s) during operation as intended, and conforms to the General Safety and Performance Requirements. |  | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
7 | POST MARKET SURVEILLANCE | [13 empty] | Section Header
| Post market documentation | PMS |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-10-29 | 2018-10-29 |  | PMS
8 | EU DECLARATION OF CONFORMITY | [13 empty] | Section Header
|  | RA | [5 empty] | Working |  |  |  | 2019-02-15 | 2019-02-15 | After QA, RA review | RA
9 | CERTIFICATES | [13 empty] | Section Header
9.1 | EU Certificate | RA | N/A | XXXX | Certificates are required for MDR compliance |  |  | Working | [6 empty] | OOS
9.2 | EN ISO 13485 Certificate | RA, QA | N/A | XXXX |  |  |  | Working | [6 empty] | OOS

## EL5ML (hidden)
LIGAMAX-5 5MM Endoscopic Multiple Clip Applier (Class IIB)
Sr. No. | Document | Responsible Function |  |  | Comments / Remarks | Deliverables | Dependency (Source of Information) | Status | Remarks | Duration (day) | Start Date | Dates without Dependency | Completion date |  | Deliverable
|  |  | From (Current Technical File TF N000006) | To (As per MDR 2017/745)
1 | DEVICE DESCRIPTION AND SPECIFICATION, INCLUDING VARIANTS AND ACCESSORIES | [13 empty] | Deliverable
1.1 | Description of Device or Device Family | [13 empty] | Section Header
1.1.1 | Product or Product Family Name | RA |  |  |  | N/A | IFU - RMC8334345, SCN053844 (P40209P12) A001036P00 | Working | Labelling team to finalize information with other functional SMEs |  |  |  | 2018-12-07 |  | RA
1.1.2 | Description of the Device or Device Family | RA |  |  |  | N/A | IFU - RMC8334345, SCN053844 (P40209P12) A001036P00 | Working | Labelling team to finalize information with other functional SMEs |  |  |  | 2018-12-07 |  | RA
1.1.3 | Intended use/Purpose | RA, Medical Affairs |  |  |  | N/A | IFU - RMC8334345, SCN053844 (P40209P12) A001036P00 | Working | Labelling team to finalize information with other functional SMEs |  |  |  | 2018-12-07 |  | RA
1.1.4 | Intended Users | RA |  |  |  | N/A | IFU - RMC8334345, SCN053844 (P40209P12) A001036P00 | Working | Labelling team to finalize information with other functional SMEs |  |  |  | 2018-12-07 |  | RA
1.1.5 | Product Codes/ Basic UDI-DI | RA |  |  | Suggested Source : GSI database GTIN | GTIN Number | N/A | Pending Review with SME |  |  |  | 2018-08-10 | 2018-12-14 |  | RA
1.1.6 | Intended Patient Population and Medical Conditions | RA, Med Affairs, QE |  |  |  | N/A | IFU - RMC8334345, SCN053844 (P40209P12) A001036P00 | Pending Information | Current IFU do not list intended Patient Population and Medical Conditions information Need to update |  |  |  | 2018-12-07 |  | Med Ops
1.1.7 | Indications and Contra-Indications | RA |  |  |  | N/A | IFU - RMC8334345, SCN053844 (P40209P12) A001036P00 | Working | Labelling team to finalize information with other functional SMEs |  |  |  | 2018-12-07 |  | RA
1.1.8 | Warnings | RA |  |  |  | N/A | IFU - RMC8334345, SCN053844 (P40209P12) A001036P00 | Working | Labelling team to finalize information with other functional SMEs |  |  |  | 2018-12-07 |  | RA
1.1.9 | Principles of Operation | RA, R&D |  |  |  | N/A | N/A | Pending Review with SME |  |  |  | 2018-12-14 | 2018-12-14 |  | R&D Elements
1.1.10 | Rationale for the Qualification of the Product as a Device | RA |  |  |  | N/A | N/A | Pending Review with SME |  |  |  | 2018-12-14 | 2018-12-14 |  | RA
1.1.11 | Device Classification and Classification rationale | RA |  |  |  | N/A | N/A | Pending Review with SME |  |  |  | 2018-12-14 | 2018-12-14 |  | RA
1.2 | Technical Specification | [13 empty] | Section Header
1.2.1 | Product Specification: | [13 empty] | Section Header
| Material Specification | R&D, Med Operations, Med Affairs, Clinical |  |  |  | Material Specification - SPE003637, SPE003638 | N/A | TD Deliverables Released |  |  |  |  | Complete |  | R&D Elements
| Design Specification | R&D, Med Operations, Med Affairs, Clinical |  |  |  | Design Specification - DR000480, DR000483 | IFU - RMC8334345, SCN053844 (P40209P12) A001036P00 | Working | ●Revise New Design Specification based on latest IFU content for Sterilization ●New Version IFUs not yet released |  |  |  | 2018-12-12 |  | R&D Elements
| Device Assembly Drawing | R&D, Med Operations, Med Affairs, Clinical |  |  |  | Drawings - LC-D44670, D44671, D44672, D44673,D44674, D44675 LC2010Z-D43237, LC3010Z - D43232, LC4010Z - D43239 | N/A | Pending Release | Need confirmation to revise current Epicentrer drawings with actual LC CAD models instead of referring to supplier drawings in EES template |  |  |  | 2018-12-12 |  | R&D Elements
1.2.2 | Novel Features | R&D, Med Operations, Med Affairs, Clinical |  |  |  | N/A | N/A | Pending Review with SME |  |  |  | 2018-08-16 | 2018-12-14 |  | R&D Elements
1.2.3 | Description of the Accessories / Combinations | R&D |  |  |  | N/A | IFU - RMC8334345, SCN053844 (P40209P12) A001036P00 | Working | Current IFUs for LC2010Z, LC3010Z and LC4010Z do not provide complete list of Product codes for Trocars that are used as accessory/ancillary Labelling team to finalize information with other functional SMEs |  |  | 2018-08-03 | 2018-12-07 |  | R&D Elements
1.2.4 | Description of Various Configurations / Variants | R&D |  |  |  | N/A | N/A | Pending Review with SME |  |  |  | 2018-08-21 | 2018-12-14 |  | R&D Elements
1.2.5 | Description of Key Functional Elements | R&D | N/A | New Section: Content in TD to be finalized with KFE Memo (SCNXXXXX) | Guidance/description : A description of the key functional elements, e.g. the device parts/components the formulation, composition, functionality and, where relevant, qualitative and quantitative composition (i.e. what parts and what material is each part made of). Where appropriate, this shall include labelled pictorial representations (e.g. diagrams, photographs, and drawings), clearly indicating key components, including sufficient explanation to understand the drawings and diagrams. | Key Functional Element Memo - SCNXXXXX | DBAM - SCN052455, SCN049258 Design Specification - DR000480, DR000483 Sterilization Strategy - SCN049267 FBD Report - PRCXXXXXX | Working | ●DBAM to be updated based on latest Biocompatibility Testing standards ●Revise New Design Specification based on latest IFU content for Sterilization ●FBD analysis report to be released after review |  |  | 2018-09-07 | 2018-12-12 |  | R&D Elements
1.2.6 | Description of Raw Materials of Key Functional Elements | R&D, Biocompatibility |  |  |  | Key Functional Element Memo - SCNXXXXX | DBAM - SCN052455, SCN049258 | Working | ●DBAM to be updated based on latest Biocompatibility Testing standards ●KFE to be released after finalizaton of other sections within the document |  |  | 2018-09-12 | 2018-12-12 | Restricted substance | R&D Elements
1.3 | Overview of Previous and Similar Generations of the Device | [13 empty] | Section Header
1.3.1 | Overview of Previous and Similar Generations of the Device | RA, R&D, Medical Operations, Clinical |  |  |  | N/A | CER - SCN052449 | Pending Information | May need updates based on latest CER after MDR related updates to CER |  |  | 2018-09-30 | 2018-12-17 |  | Med Ops
1.3.2 | Overview of Identified Similar Devices Available on the Markets | RA, Medical Operations, Clinical |  |  |  | N/A | CER - SCN052449 | Pending Information | May need updates based on latest CER after MDR related updates to CER |  |  | 2018-09-30 | 2018-12-17 |  | Med Ops
2 | INFORMATION TO BE SUPPLIED BY THE MANUFACTURER | [13 empty] | Section Header
2.1 | Labeling and Packaging Information (Inner / Outer Packaging) | Labeling, Packaging, RA |  |  |  | Package Asembly Drawing - P000506, P000507, P000413 | Explained in remarks column | Pending Information | May need updates based on latest MDR related updates to label information |  |  | 2018-11-19 | 2018-12-07 |  | LBL-LBL
2.2 | Instructions for use | Labeling, RA |  |  |  | IFU - RMC8334345, SCN053844 (P40209P12) A001036P00 | Explained in remarks column | Pending Information | May need updates based on latest MDR related updates to include information for referred TD sections : 1.1.4 1.1.6 and 1.2.3 |  |  | 2018-11-19 | 2018-12-07 |  | LBL-IFU
3 | DESIGN AND MANUFACTURING INFORMATION | [13 empty] | Section Header
3.1 | Design Stages | R&D, QE, Manufacturing Engineering |  |  |  | Design Specification - DR000480, DR000483 Material Specification - SPE003637, SPE003638 Design Verification Marix -DR000481, DR000484 Justification for Design Verification and Validation Memo -SCNXXXXXX | Explained in remarks column | Working | Design Specification - DR000444,DR000442 Justification Memo - SCNXXXXXX ●Revise New Design Specification based on latest IFU content for Sterilization ● Updated Justification Memo to be released |  |  | 2018-09-30 | 2018-12-13 |  | R&D Elements
3.2 | Manufacturing Process Information | R&D, QE, Manufacturing Engineering |  |  |  | Device Process Specification -SCN052445, SCN049259 Detailed Manufacturing Information - SCN052421, SCN052422 | Explained in remarks column | Pending Release | Updates made to the Detailed Manufacturing memo and proces specificatins to be released |  |  |  | 2018-12-07 | Clarification on MVP | Mfg Eng
3.3 | Sites including Suppliers and Sub-Contractors | R&D, RA, QE, Manufacturing Engineering, Supplier Quality |  |  |  | N/A | N/A | Pending Review with SME |  |  |  | 2018-09-30 | 2018-09-30 |  | Mfg Eng
4 | GENERAL SAFETY AND PERFORMANCE REQUIREMENTS | [13 empty] | Section Header
4.1 | General Safety and Performance Requirements Checklist | RA |  |  |  | GSPR Checklist | Explained in remarks column | Pending Review with SME | Depends on all deliverables |  |  |  | 2018-12-17 |  | RA
4.2 | List of Applied standards | Standards |  |  |  | Harmonised Standards List - SCN053711 | Explained in remarks column | Pending Review with SME | All Functional groups to conform to latest standard requirements |  |  | 2018-09-30 | 2018-12-07 |  | Standards
5 | BENEFIT-RISK ANALYSIS AND RISK MANAGEMENT | [13 empty] | Section Header
5.1 | Benefit-Risk analysis | QE, Risk Mgmt. |  |  |  | N/A | CER - SCN052449 | Pending Review with SME |  |  |  |  | 2018-12-07 |  | Risk Mgmt Docs
5.2 | Risk Management Report / Result | QE, Risk Mgmt. |  |  |  | Risk Management Report -RMD001405 Process FMEA Design - SCN052457, SCN046026 Design FMEA - RMD001406, RMD001413 Application FMEA -N/A Product Use Risk Analysis - RMD001342 Process FMEA Packaging - SCN047340 Design FMEA Packaging - SCN047339 Use Specification - SCNXXXXXX UOUP Report - SCNXXXXXX Hazard Analysis Memo - SCNXXXXX | Explained in remarks column | Working | Hazard Analysis Memo - SCNXXXXX Use Specification - SCNXXXXXX UOUP Report - SCNXXXXXX Process FMEA Design - SCN052457, SCN046026 Design FMEA - RMD001406, RMD001413 Documents mentioned above to be released after SME review |  |  | 2018-09-30 | 2018-12-13 |  | Risk Mgmt Docs
6 | PRODUCT VERIFICATION AND VALIDATION | [13 empty] | Section Header
6.1 | Pre-Clinical and Clinical Data Result | [13 empty] | Section Header
6.1.1 | Pre-Clinical and Clinical Data Results as well as Design Verification results | Pre-Clinical, Clinical, R&D |  |  |  | Design Verification and Validation Justification Memo - SCNXXXXXX Design Verifiation Matrix -DR000481, DR000484 Engineering Analysis Report - PRCXXXXX | Explained in remarks column | Working | PRCXXXXX - FBD analysis report to be released Justification Memo - SCNXXXXXX Memo to identify Worst case device to be released - SCNXXXXXX Above documents to be released after review |  |  | 10/15/18 (11/30/18 if testing required) | 2018-12-13 | Testing may require more time | Clinical-Reports
6.1.2 | Biocompatibility | Biocompatibility |  |  |  | DBAM - SCN052455, SCN049258 | N/A | Working | DBAM to be updated based on latest Biocompatibility Testing standards |  |  | 2018-11-30 | 2018-11-30 |  | BioComp Testing & Reports
6.1.3 | Devices Containing CMR or Endocrine-Disrupting Substances | Product Stewardship |  |  |  | Restricted Substance Evaluation Report - DOCXXXXXX | N/A | Working |  |  |  | 2018-10-21 | 2018-10-21 | Start time depends upon availability of information from supplier | RS Updates
6.1.4 | Electrical Safety and Electromagnetic Compatibility | R&D |  |  |  | N/A | N/A | N/A |  |  |  |  | N/A |  | R&D Elements
6.1.5 | Software Verification and Validation | R&D |  |  |  | N/A | N/A | N/A |  |  |  |  | N/A |  | R&D Elements
6.1.6 | Device Lifetime | LM, R&D |  |  |  | Product Lifetime Evaluation for EU Medical Device Regulation Requirements - DOC023641 Stability Strategy - DOC022853 | N/A | TD Deliverables Released |  |  |  | Complete | Complete |  | R&D Elements
6.1.7 | Packaging Performance and Transportation | Packaging, R&D |  |  |  | Package Transit Testing - SCN049256, SCN046018 Package Performance Qualification - SCN052425, SCN047050 | N/A | Pending Review with SME |  |  |  |  | 2018-12-07 |  | Packaging
6.1.8 | Sterilization | Sterilization |  |  |  | Sterilization Strategy - SCN049267 | N/A | Working |  |  |  | 2018-08-30 | 2018-08-30 |  | Sterilization Validation/ Reports
6.1.9 | Summary of Safety and Clinical Performance | Med Affairs, Med Ops, Clinical |  |  |  | N/A required for Class III devices | N/A | N/A |  |  |  | N/A | N/A |  | SSCP
6.1.10 | Clinical Evaluation Report | Med Affairs, Med Ops, Clinical |  |  |  | Clinical Evaluation Report - SCN052449 | Explained in remarks column | Working | ●R&D to Provide information to Principle of Operation section of CER ●Update relavant information such as intended patient population medical condition; literature study report etc. as applicable ● CER should be consistent with the information as provided within IFU and Technical documentation sections where applicable Example: May need updates based on latest MDR related updates to include information for referred TD sections: 1.3.1 and 1.3.2 |  | 2018-10-29 | 2018-12-13 | 2018-12-17 |  | CER
6.1.11 | Post Market Clinical Follow-Up Plan and Report | PMS, Med Affairs, Clinical |  |  |  | Clinical Evaluation Report -SCN052449 | PMS Plan - RMD001006 PMS Report - RMD000840 | Pending Review with SME | May need updates based on latest CER after MDR related updates to CER |  |  | 2018-10-28 | 2018-12-17 |  | PMS
6.2 | Additional Information Required in Specific Cases | [13 empty] | Section Header
6.2.1 | Devices which Incorporate Medicinal Products | R&D |  |  |  | N/A | N/A | N/A |  |  |  |  | N/A |  | R&D Elements
6.2.2 | Devices which Incorporate Tissues or Cells of Human or Animal Origin or their Derivatives | Biocompatibility |  |  |  | N/A | N/A | Pending Review with SME |  |  |  |  | 2018-12-14 |  | BioComp Testing & Reports
6.2.3 | Devices that are Composed of Substances or Combinations of Substances that are Intended to be Introduced into the Human Body and that are Absorbed by or Locally Dispersed in the Human Body) | Biocompatibility |  |  |  | N/A | N/A | Pending Review with SME |  |  |  |  | 2018-12-14 |  | BioComp Testing & Reports
6.2.4 | Devices with a Measuring Function | R&D |  |  |  | N/A | N/A | N/A |  |  |  |  | N/A |  | R&D Elements
6.2.5 | Devices Connected with other Devices in order to Operate as Intended | R&D |  |  |  | N/A | N/A | Pending Review with SME |  |  |  |  | 2018-12-14 |  | R&D Elements
7 | POST MARKET SURVEILLANCE | [13 empty] | Section Header
| Post market documentation | PMS |  |  |  | PMS Plan - RMD001006 PMS Report - RMD000840 | N/A | Pending Review with SME |  |  |  | 2018-10-28 | 2018-10-28 |  | PMS
8 | EU DECLARATION OF CONFORMITY | [13 empty] | Section Header
|  | RA | [5 empty] | Working |  |  |  | 2018-12-18 | 2018-12-18 |  | RA
9 | CERTIFICATES | [13 empty] | Section Header
9.1 | EU Certificate | RA | N/A | XXXX | Certificates to be obtained for MDR compliance |  |  | Working | [6 empty] | OOS
9.2 | EN ISO 13485 Certificate | RA, QA | N/A | XXXX |  |  |  | Working | [6 empty] | OOS
2018-12-18

## ER320-420 (hidden)
Sr. No. | Document | Responsible Function |  |  | Comments / Remarks | Deliverables | Dependency (Source of Information) | Status | Remarks | Duration (day) | Start Date | Dates without Dependency | Estimated Completion date | Associated Dependencies | Deliverable
|  |  | From (Current Technical File TF N000006) | To (As per MDR 2017/745)
1 | DEVICE DESCRIPTION AND SPECIFICATION, INCLUDING VARIANTS AND ACCESSORIES | [13 empty] | Deliverable
1.1 | Description of Device or Device Family | [13 empty] | Section Header
1.1.1 | Product or Product Family Name | RA |  |  |  | N/A | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs to update current IFU as per EU MDR |  |  | 2018-09-30 | 2019-01-30 |  | RA
1.1.2 | Description of the Device or Device Family | RA |  |  |  | N/A | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs |  |  | 2018-09-30 | 2019-01-30 |  | RA
1.1.3 | Intended use/Purpose | RA, Medical Affairs |  |  |  | N/A | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs |  |  | 2019-01-30 | 2019-01-30 | Need CER alignment | RA
1.1.4 | Intended Users | RA |  |  |  | N/A | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Pending Information | Current IFU donot list intended users information in Warnings section Labelling team to finalize information with other functional SMEs |  |  | 2019-01-30 | 2019-01-30 | Need CER alignment | RA
1.1.5 | Product Codes/ Basic UDI-DI | RA |  |  | Suggested Source : GSI database GTIN | GTIN Number | N/A | Pending Information | Need information for UDI |  |  | 2018-09-30 | 2018-09-30 |  | RA
1.1.6 | Intended Patient Population and Medical Conditions | RA, Med Affairs, QE |  |  |  | N/A | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Pending Information | Current IFU do not list intended Patient Population and Medical Conditions information Need to update |  |  | 2019-01-30 | 2019-01-30 | Need CER alignment | Med Ops
1.1.7 | Indications and Contra-Indications | RA |  |  |  | N/A | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs |  |  | 2018-01-30 | 2018-01-30 | Need CER alignment | RA
1.1.8 | Warnings | RA |  |  |  | N/A | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs |  |  | 2019-01-30 | 2019-01-30 | Need CER alignment | RA
1.1.9 | Principles of Operation | RA, R&D |  |  |  | N/A | N/A | Pending Review with SME |  |  |  | 2018-10-15 | 2018-12-17 |  | R&D Elements
1.1.10 | Rationale for the Qualification of the Product as a Device | RA |  |  |  | N/A | N/A | Pending Review with SME |  |  |  | 2018-10-30 | 2018-12-17 |  | RA
1.1.11 | Device Classification and Classification rationale | RA |  |  |  | N/A | N/A | Pending Review with SME |  |  |  | 2018-10-30 | 2018-12-17 |  | RA
1.2 | Technical Specification | [13 empty] | Section Header
1.2.1 | Product Specification: | [13 empty] | Section Header
| Material Specification | R&D, Med Operations, Med Affairs, Clinical |  |  |  | PMR35 - MS8091C PMW35 - MS8091C PRR35 - MS730006C PRW35 - MS730006C PXR35 - MS00080 PXW35 - MS00080 PSX - MS00009 | N/A | TD Deliverables Released |  |  |  |  | Complete |  | R&D Elements
| Design Specification | R&D, Med Operations, Med Affairs, Clinical |  |  |  | PMR35 - FB000551 p. 33 PMW35 - FB000551 p. 33 PRR35 - FB000501 p. 22 PRW35 - FB000501 p. 22 PXR35 - DS0057 PXW35 - DS0057 PSX - DRXXXXXX |  | Pending Review with SME | PSX - DRXXXXXX to be released in EPIcenter after review |  |  | 2018-10-15 | 2018-12-07 |  | R&D Elements
| Device Assembly Drawing | R&D, Med Operations, Med Affairs, Clinical |  |  |  | PMR35 - D02930 PMW35 - D02925 PRR35 - M09652 PRW35 - M07900 PXR35 - D05700 PXW35 - D05700 PSX - M00253 | N/A | TD Deliverables Released |  |  |  |  | Complete |  | R&D Elements
1.2.2 | Novel Features | R&D, Med Operations, Med Affairs, Clinical |  |  |  | N/A | N/A | Pending Review with SME |  |  |  | 2018-11-30 | 2018-12-07 |  | R&D Elements
1.2.3 | Description of the Accessories / Combinations | R&D |  |  |  | N/A | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Pending Review with SME |  |  |  | 2015-09-30 | 2018-12-07 |  | R&D Elements
1.2.4 | Description of Various Configurations / Variants | R&D |  |  |  | N/A | N/A | Pending Review with SME |  |  |  | 2018-09-30 | 2018-12-07 |  | R&D Elements
1.2.5 | Description of Key Functional Elements | R&D | N/A | New Section: Content in TD to be finalized with KFE Memo (SCNXXXXX) | Guidance/description : A description of the key functional elements, e.g. the device parts/components the formulation, composition, functionality and, where relevant, qualitative and quantitative composition (i.e. what parts and what material is each part made of). Where appropriate, this shall include labelled pictorial representations (e.g. diagrams, photographs, and drawings), clearly indicating key components, including sufficient explanation to understand the drawings and diagrams. | Key Functional Element Memo - SCNXXXXX | DBAM - SCN018493 Rev. XXX Design Specification - PMR35 - FB000551 p. 33 PMW35 - FB000551 p. 33 PRR35 - FB000501 p. 22 PRW35 - FB000501 p. 22 PXR35 - DS0057 PXW35 - DS0057 PSX - DRXXXXXX Sterilization Strategy - SCN005386, SCN008851 | Pending Review with SME | ●DBAM to be updated based on latest test requirements |  |  | 2018-10-15 | 2018-12-07 | Need Device testing | R&D Elements
1.2.6 | Description of Raw Materials of Key Functional Elements | R&D, Biocompatibility |  |  |  | Key Functional Element Memo - SCNXXXXX | DBAM - SCN018493 Rev. XXX | Pending Review with SME | ●DBAM to be updated based on latest test requirements ●KFE to be released after finalizaton of other sections within the document |  |  | 2018-10-15 | 2018-10-15 | Need Device testing | R&D Elements
1.3 | Overview of Previous and Similar Generations of the Device | [13 empty] | Section Header
1.3.1 | Overview of Previous and Similar Generations of the Device | RA, R&D, Medical Operations, Clinical |  |  |  | N/A | CER - SCN019399 | Pending Information | Need updates based on latest CER after MDR related updates to CER |  |  | 2018-11-30 | 2019-01-30 |  | Med Ops
1.3.2 | Overview of Identified Similar Devices Available on the Markets | RA, Medical Operations, Clinical |  |  |  | N/A | CER - SCN019399 | Pending Information | Need updates based on latest CER after MDR related updates to CER |  |  | 2018-11-30 | 2019-01-30 |  | Med Ops
2 | INFORMATION TO BE SUPPLIED BY THE MANUFACTURER | [13 empty] | Section Header
2.1 | Labeling and Packaging Information (Inner / Outer Packaging) | Labeling, Packaging, RA |  |  |  | Package Assembly Drawings PMR35- P70403 PMW35- P70527 PRR35- P70565 PRW35- P70565 PXR35- P71183 PXW35-P71183 PSX- P70459 Generic Carton Drawings PMR35- A85074 PMW35- A85074 PRR35- A85073 PRW35- A85073 PXR35- A85075 PXW35- A85075 PSX- A84734 | Explained in remarks column | Pending Information | May need updates based on latest MDR related updates to label information |  |  | 2018-11-01 | 2019-01-30 |  | LBL-LBL
2.2 | Instructions for use | Labeling, RA |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Explained in remarks column | Pending Information | May need updates based on latest MDR related updates to include information for referred TD sections : 1.1.4 1.1.6 and 1.2.3 |  |  | 2018-11-01 | 2019-01-30 |  | LBL-IFU
3 | DESIGN AND MANUFACTURING INFORMATION | [13 empty] | Section Header
3.1 | Design Stages | R&D, QE, Manufacturing Engineering |  |  |  | Factbook - FB000501, FB000551, FB001190 Design Specification - Documents mentioned above in design specification section Material Specification - Documents mentioned above in design specification section Design Verification Marix -PMR35,PMW35 - FB000551 p. 23 PRR35,PRW35 - FB000501 p. 22 PXR35,PXW35 - QA0039 | Explained in remarks column | Working | Design Specification - PSX will be released after review Design Verification Matrix - PMR and PMW is being created and will be released after review ● Updated Justification Memo to be released |  |  | 2018-10-30 | 2018-12-07 |  | R&D Elements
3.2 | Manufacturing Process Information | R&D, QE, Manufacturing Engineering |  |  |  | Manufacturing Memo for Skin Staplers and Extractor - SCNXXXXXXX | Explained in remarks column | Pending Review with SME | Manufacturing Memo for Skin Staplers and Extractor - SCNXXXXXXX to be released after review with SME |  |  | 2018-11-20 | 2018-12-07 |  | Mfg Eng
3.3 | Sites including Suppliers and Sub-Contractors | R&D, RA, QE, Manufacturing Engineering, Supplier Quality |  |  |  | N/A | N/A | Pending Review with SME |  |  |  | 2018-09-30 | 2018-12-17 |  | Mfg Eng
4 | GENERAL SAFETY AND PERFORMANCE REQUIREMENTS | [13 empty] | Section Header
4.1 | General Safety and Performance Requirements Checklist | RA |  |  |  | GSPR Checklist | Explained in remarks column | Pending Review with SME | Depends on all deliverables |  |  | 2018-11-30 | 2019-01-30 |  | RA
4.2 | List of Applied standards | Standards |  |  |  | Harmonised Standards List - SCN018477 Rev. XXX | Explained in remarks column | Pending Review with SME | All Functional groups to conform to latest standard requirements |  |  | 2018-10-30 | 2018-12-17 |  | Standards
5 | BENEFIT-RISK ANALYSIS AND RISK MANAGEMENT | [13 empty] | Section Header
5.1 | Benefit-Risk analysis | QE, Risk Mgmt. |  |  |  | CER - SCN019399 | N/A | Pending Review with SME | May need updates based on latest CER after MDR related updates to CER |  |  | 2018-11-15 | 2019-01-30 |  | Risk Mgmt Docs
5.2 | Risk Management Report / Result | QE, Risk Mgmt. |  |  |  | RMR PRR35,PRW35, PXR35,PXW35, PMR35,PMW35, PSX - RMD000284 Process FMEA (device) PRR35, PRW35, PXR35, PXW35 - PF0114 PSX - PF0270 Process FMEA (packaging) PMR35,PMW35, PSX - PF0206 PRR35,PRW35, PXR35,PXW35 - PF0206 Design FMEA (device) PMR35,PMW35 - FB000551, pg 181 (RMD001203 - working stage in EPI) PRR35,PRW35 - RMD000284 (RMD001201 - working stage in EPI) PXR35,PXW35 - DF0120 (RMD001201 - working stage in EPI) PSX - RMD000284 (RMD001188 - working stage in EPI) Design FMEA (packaging) PMR35,PMW35 - FB000551, pg 353 PRR35,PRW35 - DF0132 PXR35,PXW35 - DF0127 PSX - RMD000978 Product Use Risk Analysis (PURA) PRR35,PRW35, PXR35,PXW35, PMR35,PMW35, PSX - RMD001099 USE Spec: SCN059222 UOUP: SCN059223 | Explained in remarks column | Working | Hazard Analysis Memo - SCNXXXXX Application FMEA - RMD001200, RMD001189 Design FMEA - DF0120, RMD001203, RMD001188 Documents mentioned above to be released after SME review |  |  | 2018-11-15 | 2018-12-07 |  | Risk Mgmt Docs
6 | PRODUCT VERIFICATION AND VALIDATION | [13 empty] | Section Header
6.1 | Pre-Clinical and Clinical Data Result | [13 empty] | Section Header
6.1.1 | Pre-Clinical and Clinical Data Results as well as Design Verification results | Pre-Clinical, Clinical, R&D |  |  |  | Design Verifiation Matrix -PMR35,PMW35 - FB000551 p. 23 PRR35,PRW35 - FB000501 p. 22 PXR35,PXW35 - QA0039 PSX - N/A Related Testing documents such as Reliability, Usability and other validation evidences etc. - Factbook information/CER | Explained in remarks column | Working | Design Verifiation Matrix - PMR and PMW to be reviewed by SME along with associated Preclinical information Above documents to be released after review |  |  | 2018-11-30 | 2018-12-17 |  | Clinical-Reports
6.1.2 | Biocompatibility | Biocompatibility |  |  |  | DBAM - SCN018493 Rev. XXX | N/A | Working | DBAM to be updated based on latest Biocompatibility Testing standards |  |  | 2019-02-15 | 2019-02-15 | Need Device testing | BioComp Testing & Reports
6.1.3 | Devices Containing CMR or Endocrine-Disrupting Substances | Product stewardship |  |  |  | Restricted Substance Evaluation Report - DOCXXXXXX | N/A | Working |  |  |  | 2018-10-15 | 2018-12-17 |  | RS Updates
6.1.4 | Electrical Safety and Electromagnetic Compatibility | R&D |  |  |  | N/A | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
6.1.5 | Software Verification and Validation | R&D |  |  |  | N/A | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
6.1.6 | Device Lifetime | LM, R&D |  |  |  | Product Lifetime Evaluation for EU Medical Device Regulation Requirements - DOC023641 Stability Strategy (device) PMR35,PMW35 - FB000551, p221 PRR35,PRW35 - DOC022617 PXR35,PXW35 - FB001190, p 418 PSX - DOC022518 Stability Strategy (Package) PRR35,PRW35, PXR35,PXW35, PMR35,PMW35, PSX - SCN003465 | N/A | TD Deliverables Released |  |  |  | 2018-09-15 | Complete |  | R&D Elements
6.1.7 | Packaging Performance and Transportation | Packaging, R&D |  |  |  | Package Validation (Transit test) PMR35,PMW35 - FB000551 p. 538 PRR35,PRW35 - PRC037643 PXR35,PXW35 - PRC037646 PSX - PRC003331 | N/A | TD Deliverables Released |  |  |  | 2018-10-31 | Complete |  | Packaging
6.1.8 | Sterilization | Sterilization |  |  |  | Sterilization strategy PRR35,PRW35, PXR35,PXW35, PMR35,PMW35 - SCN005386 Rev. XXX PSX - SCN008851 Rev. XXX | N/A | Pending Review with SME |  |  |  | 2018-10-30 | 2018-12-17 |  | Sterilization Validation/ Reports
6.1.9 | Summary of Safety and Clinical Performance | Med Affairs, Med Ops, Clinical |  |  |  | N/A required for Class III devices | N/A | N/A |  |  |  | 2018-10-30 | N/A |  | SSCP
6.1.10 | Clinical Evaluation Report | Med Affairs, Med Ops, Clinical |  |  |  | Clinical Evaluation Report - SCN019399 | Explained in remarks column | Working | ●R&D to Provide information to Principle of Operation section of CER ●Update relavant information such as intended patient population medical condition; literature study report etc. as applicable ● CER should be consistent with the information as provided within IFU and Technical documentation sections where applicable Example: May need updates based on latest MDR related updates to include information for referred TD sections: 1.3.1 and 1.3.2 |  |  | 2019-01-21 | 2019-01-21 | Need Biocompatibility input | CER
6.1.11 | Post Market Clinical Follow-Up Plan and Report | PMS, Med Affairs, Clinical |  |  |  | Clinical Evaluation Report - SCN019399 | PMS Plan - RMD001036 PMS Reports data for CER - RMD000830 | Pending Review with SME | May need updates based on latest CER after MDR related updates to CER |  |  | 2018-10-29 | 2019-01-21 |  | PMS
6.2 | Additional Information Required in Specific Cases | [13 empty] | Section Header
6.2.1 | Devices which Incorporate Medicinal Products | R&D |  |  |  | N/A | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
6.2.2 | Devices which Incorporate Tissues or Cells of Human or Animal Origin or their Derivatives | Biocompatibility |  |  |  | N/A | N/A | Pending Review with SME |  |  |  | 2018-08-30 | 2018-08-30 |  | BioComp Testing & Reports
6.2.3 | Devices that are Composed of Substances or Combinations of Substances that are Intended to be Introduced into the Human Body and that are Absorbed by or Locally Dispersed in the Human Body) | Biocompatibility |  |  |  | N/A | N/A | Pending Review with SME |  |  |  | 2018-08-30 | 2018-08-30 |  | BioComp Testing & Reports
6.2.4 | Devices with a Measuring Function | R&D |  |  |  | N/A | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
6.2.5 | Devices Connected with other Devices in order to Operate as Intended | R&D |  |  |  | N/A | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
7 | POST MARKET SURVEILLANCE | [13 empty] | Section Header
| Post market documentation | PMS |  |  |  | PMS Plan - RMD001036 Rev. XXX PMS Report - RMD000830 Rev. XXX PSUR - TBD | N/A | Pending Review with SME |  |  |  | 2018-10-29 | 2018-10-29 |  | PMS
8 | EU DECLARATION OF CONFORMITY | [13 empty] | Section Header
|  | RA | [5 empty] | Working |  |  |  | 2019-02-15 | 2019-02-15 | After QA, RA review | RA
9 | CERTIFICATES | [13 empty] | Section Header
9.1 | EU Certificate | RA | N/A | XXXX | Certificates are required for MDR compliance |  |  | Working | [6 empty] | OOS
9.2 | EN ISO 13485 Certificate | RA, QA | N/A | XXXX |  |  |  | Working | [6 empty] | OOS

## MCA (hidden)
Sr. No. | Document | Responsible Function |  |  | Comments / Remarks | Deliverables | Dependency (Source of Information) | Status | Remarks | Duration (day) | Start Date | Dates without Dependency | Estimated Completion date | Associated Dependencies | Deliverable
|  |  | From (Current Technical File TF N000006) | To (As per MDR 2017/745)
1 | DEVICE DESCRIPTION AND SPECIFICATION, INCLUDING VARIANTS AND ACCESSORIES | [13 empty] | Deliverable
1.1 | Description of Device or Device Family | [13 empty] | Section Header
1.1.1 | Product or Product Family Name | RA |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs to update current IFU as per EU MDR |  |  | 2018-09-30 | 2019-01-30 |  | RA
1.1.2 | Description of the Device or Device Family | RA |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs |  |  | 2018-09-30 | 2019-01-30 |  | RA
1.1.3 | Intended use/Purpose | RA, Medical Affairs |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs |  |  | 2019-01-30 | 2019-01-30 | Need CER alignment | RA
1.1.4 | Intended Users | RA |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Pending Information | Current IFU donot list intended users information in Warnings section Labelling team to finalize information with other functional SMEs |  |  | 2019-01-30 | 2019-01-30 | Need CER alignment | RA
1.1.5 | Product Codes/ Basic UDI-DI | RA |  |  | Suggested Source : GSI database GTIN |  | N/A | Pending Information | Need information for UDI |  |  | 2018-09-30 | 2018-09-30 |  | RA
1.1.6 | Intended Patient Population and Medical Conditions | RA, Med Affairs, QE |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Pending Information | Current IFU do not list intended Patient Population and Medical Conditions information Need to update |  |  | 2019-01-30 | 2019-01-30 | Need CER alignment | Med Ops
1.1.7 | Indications and Contra-Indications | RA |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs |  |  | 2018-01-30 | 2018-01-30 | Need CER alignment | RA
1.1.8 | Warnings | RA |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Working | Labelling team to finalize information with other functional SMEs |  |  | 2019-01-30 | 2019-01-30 | Need CER alignment | RA
1.1.9 | Principles of Operation | RA, R&D |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-10-15 | 2018-12-17 |  | R&D Elements
1.1.10 | Rationale for the Qualification of the Product as a Device | RA |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-10-30 | 2018-12-17 |  | RA
1.1.11 | Device Classification and Classification rationale | RA |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-10-30 | 2018-12-17 |  | RA
1.2 | Technical Specification | [13 empty] | Section Header
1.2.1 | Product Specification: | [13 empty] | Section Header
| Material Specification | R&D, Med Operations, Med Affairs, Clinical |  |  |  |  | N/A | TD Deliverables Released |  |  |  |  | Complete |  | R&D Elements
| Design Specification | R&D, Med Operations, Med Affairs, Clinical | [5 empty] | Pending Review with SME | PSX - DRXXXXXX to be released in EPIcenter after review |  |  | 2018-10-15 | 2018-12-07 |  | R&D Elements
| Device Assembly Drawing | R&D, Med Operations, Med Affairs, Clinical |  |  |  |  | N/A | TD Deliverables Released |  |  |  |  | Complete |  | R&D Elements
1.2.2 | Novel Features | R&D, Med Operations, Med Affairs, Clinical |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-11-30 | 2018-12-07 |  | R&D Elements
1.2.3 | Description of the Accessories / Combinations | R&D |  |  |  |  | IFU- A88260P00, A86107P00, A88261P00, A88259P00 | Pending Review with SME |  |  |  | 2015-09-30 | 2018-12-07 |  | R&D Elements
1.2.4 | Description of Various Configurations / Variants | R&D |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-09-30 | 2018-12-07 |  | R&D Elements
1.2.5 | Description of Key Functional Elements | R&D | N/A | New Section: Content in TD to be finalized with KFE Memo (SCNXXXXX) | Guidance/description : A description of the key functional elements, e.g. the device parts/components the formulation, composition, functionality and, where relevant, qualitative and quantitative composition (i.e. what parts and what material is each part made of). Where appropriate, this shall include labelled pictorial representations (e.g. diagrams, photographs, and drawings), clearly indicating key components, including sufficient explanation to understand the drawings and diagrams. |  | DBAM - SCN018493 Rev. XXX Design Specification - PMR35 - FB000551 p. 33 PMW35 - FB000551 p. 33 PRR35 - FB000501 p. 22 PRW35 - FB000501 p. 22 PXR35 - DS0057 PXW35 - DS0057 PSX - DRXXXXXX Sterilization Strategy - SCN005386, SCN008851 | Pending Review with SME | ●DBAM to be updated based on latest test requirements |  |  | 2018-10-15 | 2018-12-07 | Need Device testing | R&D Elements
1.2.6 | Description of Raw Materials of Key Functional Elements | R&D, Biocompatibility |  |  |  |  | DBAM - SCN018493 Rev. XXX | Pending Review with SME | ●DBAM to be updated based on latest test requirements ●KFE to be released after finalizaton of other sections within the document |  |  | 2018-10-15 | 2018-10-15 | Need Device testing | R&D Elements
1.3 | Overview of Previous and Similar Generations of the Device | [13 empty] | Section Header
1.3.1 | Overview of Previous and Similar Generations of the Device | RA, R&D, Medical Operations, Clinical |  |  |  |  | CER - SCN019399 | Pending Information | Need updates based on latest CER after MDR related updates to CER |  |  | 2018-11-30 | 2019-01-30 |  | Med Ops
1.3.2 | Overview of Identified Similar Devices Available on the Markets | RA, Medical Operations, Clinical |  |  |  |  | CER - SCN019399 | Pending Information | Need updates based on latest CER after MDR related updates to CER |  |  | 2018-11-30 | 2019-01-30 |  | Med Ops
2 | INFORMATION TO BE SUPPLIED BY THE MANUFACTURER | [13 empty] | Section Header
2.1 | Labeling and Packaging Information (Inner / Outer Packaging) | Labeling, Packaging, RA |  |  |  |  | Explained in remarks column | Pending Information | May need updates based on latest MDR related updates to label information |  |  | 2018-11-01 | 2019-01-30 |  | LBL-LBL
2.2 | Instructions for use | Labeling, RA |  |  |  |  | Explained in remarks column | Pending Information | May need updates based on latest MDR related updates to include information for referred TD sections : 1.1.4 1.1.6 and 1.2.3 |  |  | 2018-11-01 | 2019-01-30 |  | LBL-IFU
3 | DESIGN AND MANUFACTURING INFORMATION | [13 empty] | Section Header
3.1 | Design Stages | R&D, QE, Manufacturing Engineering |  |  |  |  | Explained in remarks column | Working | Design Specification - PSX will be released after review Design Verification Matrix - PMR and PMW is being created and will be released after review ● Updated Justification Memo to be released |  |  | 2018-10-30 | 2018-12-07 |  | R&D Elements
3.2 | Manufacturing Process Information | R&D, QE, Manufacturing Engineering |  |  |  |  | Explained in remarks column | Pending Review with SME | Manufacturing Memo for Skin Staplers and Extractor - SCNXXXXXXX to be released after review with SME |  |  | 2018-11-20 | 2018-12-07 |  | Mfg Eng
3.3 | Sites including Suppliers and Sub-Contractors | R&D, RA, QE, Manufacturing Engineering, Supplier Quality |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-09-30 | 2018-12-17 |  | Mfg Eng
4 | GENERAL SAFETY AND PERFORMANCE REQUIREMENTS | [13 empty] | Section Header
4.1 | General Safety and Performance Requirements Checklist | RA |  |  |  |  | Explained in remarks column | Pending Review with SME | Depends on all deliverables |  |  | 2018-11-30 | 2019-01-30 |  | RA
4.2 | List of Applied standards | Standards |  |  |  |  | Explained in remarks column | Pending Review with SME | All Functional groups to conform to latest standard requirements |  |  | 2018-10-30 | 2018-12-17 |  | Standards
5 | BENEFIT-RISK ANALYSIS AND RISK MANAGEMENT | [13 empty] | Section Header
5.1 | Benefit-Risk analysis | QE, Risk Mgmt. |  |  |  |  | N/A | Pending Review with SME | May need updates based on latest CER after MDR related updates to CER |  |  | 2018-11-15 | 2019-01-30 |  | Risk Mgmt Docs
5.2 | Risk Management Report / Result | QE, Risk Mgmt. |  |  |  |  | Explained in remarks column | Working | Hazard Analysis Memo - SCNXXXXX Application FMEA - RMD001200, RMD001189 Design FMEA - DF0120, RMD001203, RMD001188 Documents mentioned above to be released after SME review |  |  | 2018-11-15 | 2018-12-07 |  | Risk Mgmt Docs
6 | PRODUCT VERIFICATION AND VALIDATION | [13 empty] | Section Header
6.1 | Pre-Clinical and Clinical Data Result | [13 empty] | Section Header
6.1.1 | Pre-Clinical and Clinical Data Results as well as Design Verification results | Pre-Clinical, Clinical, R&D |  |  |  |  | Explained in remarks column | Working | Design Verifiation Matrix - PMR and PMW to be reviewed by SME along with associated Preclinical information Above documents to be released after review |  |  | 2018-11-30 | 2018-12-17 |  | Clinical-Reports
6.1.2 | Biocompatibility | Biocompatibility |  |  |  |  | N/A | Working | DBAM to be updated based on latest Biocompatibility Testing standards |  |  | 2019-02-15 | 2019-02-15 | Need Device testing | BioComp Testing & Reports
6.1.3 | Devices Containing CMR or Endocrine-Disrupting Substances | Product stewardship |  |  |  |  | N/A | Working |  |  |  | 2018-10-15 | 2018-12-17 |  | RS Updates
6.1.4 | Electrical Safety and Electromagnetic Compatibility | R&D |  |  |  |  | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
6.1.5 | Software Verification and Validation | R&D |  |  |  |  | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
6.1.6 | Device Lifetime | LM, R&D |  |  |  |  | N/A | TD Deliverables Released |  |  |  | 2018-09-15 | Complete |  | R&D Elements
6.1.7 | Packaging Performance and Transportation | Packaging, R&D |  |  |  |  | N/A | TD Deliverables Released |  |  |  | 2018-10-31 | Complete |  | Packaging
6.1.8 | Sterilization | Sterilization |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-10-30 | 2018-12-17 |  | Sterilization Validation/ Reports
6.1.9 | Summary of Safety and Clinical Performance | Med Affairs, Med Ops, Clinical |  |  |  |  | N/A | N/A |  |  |  | 2018-10-30 | N/A |  | SSCP
6.1.10 | Clinical Evaluation Report | Med Affairs, Med Ops, Clinical |  |  |  |  | Explained in remarks column | Working | ●R&D to Provide information to Principle of Operation section of CER ●Update relavant information such as intended patient population medical condition; literature study report etc. as applicable ● CER should be consistent with the information as provided within IFU and Technical documentation sections where applicable Example: May need updates based on latest MDR related updates to include information for referred TD sections: 1.3.1 and 1.3.2 |  |  | 2019-01-21 | 2019-01-21 | Need Biocompatibility input | CER
6.1.11 | Post Market Clinical Follow-Up Plan and Report | PMS, Med Affairs, Clinical |  |  |  |  | PMS Plan - RMD001036 PMS Reports data for CER - RMD000830 | Pending Review with SME | May need updates based on latest CER after MDR related updates to CER |  |  | 2018-10-29 | 2019-01-21 |  | PMS
6.2 | Additional Information Required in Specific Cases | [13 empty] | Section Header
6.2.1 | Devices which Incorporate Medicinal Products | R&D |  |  |  |  | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
6.2.2 | Devices which Incorporate Tissues or Cells of Human or Animal Origin or their Derivatives | Biocompatibility |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-08-30 | 2018-08-30 |  | BioComp Testing & Reports
6.2.3 | Devices that are Composed of Substances or Combinations of Substances that are Intended to be Introduced into the Human Body and that are Absorbed by or Locally Dispersed in the Human Body) | Biocompatibility |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-08-30 | 2018-08-30 |  | BioComp Testing & Reports
6.2.4 | Devices with a Measuring Function | R&D |  |  |  |  | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
6.2.5 | Devices Connected with other Devices in order to Operate as Intended | R&D |  |  |  |  | N/A | N/A |  |  |  | 2018-08-30 | N/A |  | R&D Elements
7 | POST MARKET SURVEILLANCE | [13 empty] | Section Header
| Post market documentation | PMS |  |  |  |  | N/A | Pending Review with SME |  |  |  | 2018-10-29 | 2018-10-29 |  | PMS
8 | EU DECLARATION OF CONFORMITY | [13 empty] | Section Header
|  | RA | [5 empty] | Working |  |  |  | 2019-02-15 | 2019-02-15 | After QA, RA review | RA
9 | CERTIFICATES | [13 empty] | Section Header
9.1 | EU Certificate | RA | N/A | XXXX | Certificates are required for MDR compliance |  |  | Working | [6 empty] | OOS
9.2 | EN ISO 13485 Certificate | RA, QA | N/A | XXXX |  |  |  | Working | [6 empty] | OOS

## Dimension & material specs
All-In-One Hand Control Unit and Electrodes
| Model | Description | Instructions for use | Delivery status | Material of the body-contact part | Dimension | Outer diameter of the inserted portion
Before change | 0080 | Electrode (L-Wire Hook with suction function laparoscopic non-stick electrode, 38 cm) | Disposable | Sterile | PTFE coating | 38 cm | 2 mm
After change | 0080J |  | Disposable | Sterile | PTFE coating | 38 cm | 2 mm
Comments:
No changes in dimensionsal and material specifications

## Requirements and Test methods
Section | Draft YY Requirement/Testing | Test method/Supporting Evidence | Supporting Evidence if any test conducted
3.1.1 | Appearance
| No surface defect(rust,scar) | Visual inspection as per TRD
3.1.2 | Dimensions
| Tolerance of ± 10% | standard measuring tools to be used as per TRD
3.1.3 | Connection firmness
| The connection between the electrode and the electrosurgical pencil should be firm and able to withstand a force of 15 N without looseness. | Apply a force of 15 N between the electrode and the electrosurgical pencil for 15 s; the result shall comply with the requirements (Test procedure is defined in TRD)
3.1.4 | Corossion Resistance Test
| the corrosion of the electrode should at least meet the requirements of Grade b of the boiling water method specified in YY/T 0149-2006 | Boiling water method (YY/T 0149-2006 document available and test procedure defined in doc)
3.1.5 | Outer Diameter
| < 2 mm(lower than) GB 11244-2005 | standard measuring tools to be used as per TRD (document not available)
3.3 | Suction flow rate
| >500 mL/min(greater than) | Operate the machine to run stably. The pressure of the vacuum suction device should not be lower than 0.09 Mpa. The standard measuring cylinder should be used for measurement. The result should comply with the requirements (Test procedure is defined in TRD)
3.4 | Electromagnetic compatibility
| It shall comply with the requirements of YY 0505-2012 and Clause 36 of GB 9706.4-2009 | (YY 0505-2012 and Clause 36 of GB 9706.4-2009 document available and test procedure defined in doc) |  | m
Comments: 1.Electrical safety test are to be conducted as per GB 9706.1-2007, GB 9706.4-2009 and GB 9706.19-2000 (GB documents available, Types of test to be conducted are not defined in TRD)

## Packaging and identification
Clause No. | Draft YY Requirement/Testing | Supporting Evidence
Packaging
8 | Identification for single package
Comments 1.Documents not received

## Variants
Model | Description | Instructions for use | Delivery status | Material of the body-contact part | Dimension | Outer diameter of the inserted portion
All-In-One Hand Control Unit and Electrodes
55 | All-In-One Hand Piece | Reusable | Non-sterile | N/A | 23.5 cm | N/A
60 | Electrode (L-Hook with suction function laparoscopic non-stick electrode, 32 cm) | Disposable | Sterile | PTFE coating | 32 cm | 2 mm
85 | Electrode (J-Wire Hook with suction function laparoscopic non-stick electrode, 38 cm) | Disposable | Sterile | PTFE coating | 38 cm | 2 mm
62 | Electrode (Curved Blade with suction function laparoscopic non-stick electrode, 32 cm) | Disposable | Sterile | PTFE coating | 32 cm | 2 mm
63 | Electrode (L-Wire Hook with suction function laparoscopic non-stick electrode, 32 cm) | Disposable | Sterile | PTFE coating | 32 cm | 2 mm
80 | Electrode (L-Wire Hook with suction function laparoscopic non-stick electrode, 38 cm) | Disposable | Sterile | PTFE coating | 38 cm | 2 mm
64 | Electrode (Standard Blade with suction function laparoscopic non-stick electrode, 32 cm) | Disposable | Sterile | PTFE coating | 32 cm | 2 mm
Laparoscopic electrodes
17 | Electrode (Standard Blade laparoscopic non-stick electrode, 13.5 inch) | Disposable | Sterile | PTFE coating | 33 cm | 2 mm
18 | Electrode (Spatula laparoscopic non-stick electrode, 13.5 inch) | Disposable | Sterile | PTFE coating | 33 cm | 2 mm
0018C | Electrode (Curved Spatula laparoscopic non-stick electrode, 13.5 inch) | Disposable | Sterile | PTFE coating | 33 cm | 2 mm
0018CS | Electrode (Curved Spatula laparoscopic non-stick electrode/SPLIT STEM, 13.5 inch) | Disposable | Sterile | PTFE coating | 33 cm | 3.5 cm
19 | Electrode (Curved Blade laparoscopic non-stick electrode, 13.5 inch) | Disposable | Sterile | PTFE coating | 33 cm | 2 mm
0019L | Electrode (Curved Blade laparoscopic non-stick electrode, 45 cm) | Disposable | Sterile | PTFE coating | 45 cm | 2 mm
0019S | Electrode (Curved Blade laparoscopic non-stick electrode/SPLIT STEM, 13.5 inch) | Disposable | Sterile | PTFE coating | 33 cm | 3.5 cm
0019LS | Electrode (Curved Blade laparoscopic non-stick electrode/SPLIT STEM, 45 cm) | Disposable | Sterile | PTFE coating | 45 cm | 3.5 cm
20 | Electrode (L-Hook laparoscopic non-stick electrode, 13.5 inch) | Disposable | Sterile | PTFE coating | 33 cm | 2 mm
0020L | Electrode (L-Hook laparoscopic non-stick electrode, 45 cm) | Disposable | Sterile | PTFE coating | 45 cm | 2 mm
0020S | Electrode (L-Hook laparoscopic non-stick electrode/SPLIT STEM, 13.5 inch) | Disposable | Sterile | PTFE coating | 33 cm | 3.5 cm
0020LS | Electrode (L-Hook laparoscopic non-stick electrode/SPLIT STEM, 45 cm) | Disposable | Sterile | PTFE coating | 45 cm | 3.5 cm
21 | Electrode (J-Hook laparoscopic non-stick electrode, 13.5 inch) | Disposable | Sterile | PTFE coating | 33 cm | 2 mm
0021L | Electrode (J-Hook laparoscopic non-stick electrode, 45 cm) | Disposable | Sterile | PTFE coating | 45 cm | 2 mm
0021S | Electrode (J-Hook laparoscopic non-stick electrode/SPLIT STEM, 13.5 inch) | Disposable | Sterile | PTFE coating | 33 cm | 3.5 cm
0021LS | Electrode (J-Hook laparoscopic non-stick electrode/SPLIT STEM, 45 cm) | Disposable | Sterile | PTFE coating | 45 cm | 3.5 cm
22 | Electrode (Needle laparoscopic non-stick electrode) | Disposable | Sterile | PTFE coating | 33 cm | 2 mm
0022S | Electrode (Needle laparoscopic non-stick electrode/SPLIT STEM) | Disposable | Sterile | PTFE coating | 33 cm | 3.5 cm
24 | Electrode (Ball laparoscopic non-stick electrode, 13.5 inch) | Disposable | Sterile | PTFE coating | 33 cm | 2 mm
0026S | Electrode (O-Wire laparoscopic non-stick electrode/SPLIT STEM, 13.5 inch) | Disposable | Sterile | PTFE coating | 33 cm | 3.5 cm
100 | Electrode (L-Wire laparoscopic non-stick electrode, 13.5 inch) | Disposable | Sterile | PTFE coating | 33 cm | 2 mm
0100L | Electrode (L-Wire laparoscopic non-stick electrode, 45 cm) | Disposable | Sterile | PTFE coating | 45 cm | 2 mm
0100S | Electrode (L-Wire laparoscopic non-stick electrode/SPLIT STEM, 13.5 inch) | Disposable | Sterile | PTFE coating | 33 cm | 3.5 cm
0100LS | Electrode (L-Wire laparoscopic non-stick electrode/SPLIT STEM, 45 cm) | Disposable | Sterile | PTFE coating | 45 cm | 3.5 cm
0105S | Electrode (J-Wire laparoscopic non-stick electrode/SPLIT STEM) | Disposable | Sterile | PTFE coating | 33 cm | 5.0 mm