## Product Specification&materials
Product Specification Table
Model | Electrode width | Electrode length | Sterilization method | Shelf life | Model naming
ME725M1C | 0.27cm+0.003/-0.008cm | 6.35cm | Ethylene oxide | 3 years | ME725M1C ME refers to Megadyne 7 refers to Ace blade 700 25 refers to 2.5 inches, the electrode length, which is equivalent to 6.35cm mentioned in table above M refers to modified electrodes 1C refers to a C interface
ME725M1E | 0.27cm+0.003/-0.008cm | 6.35cm | Ethylene oxide | 3 years | ME725M1E ME refers to Megadyne 7 refers to Ace blade 700 25 refers to 2.5 inches, the electrode length, which is equivalent to 6.35cm mentioned in table above M refers to modified electrodes 1E refers to an E interface
ME7251C | 0.27cm+0.003/-0.008cm | 6.35cm | Gamma ray radiation | 3 years | ME7251C ME refers to Megadyne 7 refers to Ace blade 700 25 refers to 2.5 inches, the electrode length, which is equivalent to 6.35cm mentioned in table above 1C refers to a C interface
ME7251E | 0.27cm+0.003/-0.008cm | 6.35cm | Gamma ray radiation | 3 years | ME7251E ME refers to Megadyne 7 refers to Ace blade 700 25 refers to 2.5 inches, the electrode length, which is equivalent to 6.35cm mentioned in table above 1E refers to an E interface
Materials in contact with patients
Blade | 305 stainless steel
Shaft and tip insulation | Polyolefin, PTFE
Electrode coating | PTFE (Polytetrafluoroethylene)
Pad printing | Ink, (Marabu TPU970)
End cover | Vinyl, TPE (green)
Modified electrode distal tip insulation | PTFE
Nose cone | Silicone rubber

## Product and its labelling
Schematic diagram of electrode:
|  | Standard electrode: | Modified electrode:
Component
1 Electrode | 6 Protective sleeve
2 Exhaust pipe | 7 Exhauster connector
3 Cut button | 8 Generator connector
4 Coagulation button | 9 Connecting power cord
5 Electrosurgical pencil body | 10 Gray sumitomo A4 heat-shrinkable tube

## Requirements and Test methods
Section | Draft YY Requirement/Testing | Requirements | Supporting Evidence if any test conducted
2.1 | Safety properties
2.1.1 | The general requirements for safety of surgical electrode shall comply with the requirements of GB 9706.1-2007. For the safety characteristics, see Annex A. | Conduct the test on the general requirements for safety of surgical electrode according to the method specified in GB 9706.1-2007; the result shall comply with the listed requirements.
2.1.2 | The particular requirements for safety of surgical electrode shall comply with the requirements of GB 9706.4-2009. | Conduct the test on the particular requirements for safety of surgical electrode according to the methods specified in GB 9706.4-2009; the result shall comply with the requirements of 2.1.2.
2.1.3 | The EMC requirements of surgical electrode shall comply with the requirements of YY 0505-2012 and Clause 36 in GB 9706.4-2009. | Conduct the test on the particular requirements for safety of surgical electrode according to the methods specified in YY 0505-2012 and Clause 36 in GB 9706.4-2009; the result shall comply with the requirements of 2.1.3.
2.2 | Appearance
2.2.1 | The surface of the electrosurgical pencil shall be clean and free from spots, as well as the rough surface causing damages.
2.2.2 | The surface of electrode shall be free from rusting, corrosion spot or other surface defects. | Perform a visual inspection and hand feeling test; the result shall comply with the requirements of 2.2.
2.2.3 | The surface of the exhaust pipe shall be clean without any manufacturing defects.
2.3 | Sterility
2.3.1 | Electrosurgical pencil body, electrode and exhaust pipe shall be sterilized. | Conduct the test according to the method specified in the Chinese Pharmacopoeia (Edition 2015, Volume IV); the result shall comply with the requirements of 2.3.1.
2.3.2 | Ethylene oxide residues: The EO residues of parts in contact with the patient shall not exceed 10 μg/g. | Ethylene oxide residues: Conduct the test according to the method specified in GB/T 14233.1-2008; the result shall comply with the listed requirements.
2.4 | General Properties
2.4.1 | The length of the electrosurgical pencil body is 14.43 cm, with the allowable error of ± 0.1cm. The length of electrode is 6.35 cm, with the allowable error of ± 0.03 cm. The length of the connected power cord is 3 m, with the allowable error of ±15.2 cm.
2.4.2 | Continuity of electrode: The electrode shall be well connected to the host cable and the impedance value shall be less than 50 Ω.
2.4.3 | The hardness of metal parts of electrode (component 1) shall not be less than 120HV0.2
2.4.4 | The surface roughness (Ra) of electrode (component 1) shall be not bigger than 0.5 μm.
2.4.5 | After the corrosion resistance test is completed, the corrosion degree of electrode (component 1) shall be lower than the value listed in Class B requirements for boiling water treatment specified in YY/T 0149-2006.

## Annex A
Product characteristics
a) Classification according to the degree of protection against electric shock: Not applicable
b) Classification according to the type of protection against electric shock: Type CF
c) Classification according to the degree of protection against liquid ingress: Not applicable
d) Classification according to the safety degree of product used in the flammable anesthesia gas mixed with air, oxygen or nitrous oxide: Not applicable
e) Classification according to the operation mode: Not applicable
f) Rated voltage and frequency of equipment: Not applicable
g) Input power of equipment: Not applicable
h) This equipment has no defibrillation-proof applied part;
i) This equipment has no signal output or input part;
j) Permanently installed equipment or non-permanently installed equipment: Non-permanently installed equipment;
k) Electrical insulation diagram (Not applicable, with the host)
l) Rated voltage of accessory: 5.5kVp

## Packaging and identification
Clause No. | Draft YY Requirement/Testing | Supporting Evidence
Packaging
8 | Identification for single package
Comments 1.Documents not received