|  | THIS DOCUMENT IS PROVIDED BY THE APPLICANT/THE REGISTRANT |
| --- | --- |

Serial number of product technical requirements for medical device:

Patient Return Electrodes

- Product model/specification

  - Degree of protection against electric shock: Type CF applied part.

  - Specifications/models: The product is reusable and packaged in non-sterile state; for models/specifications, see Table 1 for details.

Table 1: Table of specifications/models

| Model | English description |
| --- | --- |
| Patient Return Electrodes |  |
| 0840 | MEGA SOFT Reusable Pediatric Patient Return Electrode |
| 0830 | MEGA SOFT® - Patient Return Electrode |
| 0835 | MEGA SOFT® Dual Cord - Patient Return Electrode |
| 0800 | MEGA 2000® Reusable Patient Return Electrode |

  - Shelf life of product: 2 years

- Technical requirements

  - Appearance: The external surface of Patient Return Electrodes shall be clean and free from spots, as well as the rough surface causing damages.

  - Connector: The connector of Patient Return Electrodes shall be clear and reliable without looseness.

  - Connection firmness: The connections between connectors of Patient Return Electrodes shall be able to withstand a force of 40 N without fracture.

  - Fitting performance: Patient Return Electrodes shall be free from looseness and the combined machines have good service performance.

  - Electrical safety

    - Product safety characteristics

Degree of protection against electric shock: Type CF applied part.

    - Requirements

Patient Return Electrodes shall comply with the requirements of GB 9706.1-2007 Medical electrical equipment - Part 1: General requirements for safety and GB 9706.4-2009 Medical electrical equipment - Part 2: Particular requirements for the safety of high frequency surgical equipment. 2.6 Patient Return Electrodes shall comply with the requirements of YY 0505-2012 Medical electrical equipment - Part 1-2: General requirements for safety - Collateral standard: Electromagnetic compatibility - Requirements and tests.

- Test methods

  - Appearance: Verification shall be checked through visual inspection.

  - Connector: Verification shall be checked through inspection.

  - Connection firmness: Apply a force of 40 N to the connectors between Patient Return Electrodes for 15 s; the result shall comply with the requirements of 4.4.

  - Fitting performance: Verification shall be checked by operation of the product.

  - Electrical safety: Conduct the test according to the methods specified in GB 9706.1-2007 and GB 9706.4-2009.

  - Electromagnetic compatibility test: Conduct the test according to the methods specified in YY 0505-2012 Medical electrical equipment - Part 1-2: General requirements for safety - Collateral standard: Electromagnetic compatibility - Requirements and tests, and GB 9706.4-2009 Medical electrical equipment.