## Slide 1
Annual Quality Review Meeting
Seolin Systems
Korea Service Center Date:20-05-2020
Ethicon Endo-Surgery

## Slide 2: Attendees
Name | Title
Hwang, Min Ho | J&J Medical Korea
Kim, Kyung Ho | Seolin system (Korea Service Center)
Shannon Gillespie | International Service Manager
Robert Peters | Service Quality Lead
James Swords | WTCS-Coordinator II
James Terry | Product Release Technician
Janet Holt | WTCS-Coordinator II
Notes: Same from 2019 for Korea Team.

## Slide 3: Agenda
Organizational changes
Follow-up items from previous quality review
Significant changes in Quality System
Average monthly service volume during last 12 months
Service training results
Product quality review
Scorecard Metrics
Management Review Summary
Compliance Update
Certifications
Records Review for destruction
General discussion
Agenda

## Slide 4
Organizational Changes

## Slide 5
Organizational changes
.
CEO
(H.T Hwang)
Total : 28Person
Marketing
Department
Sales (1st)
Department
Sales (2nd)
Department
Sales (3rd)
Department
Service
Department
ETHICON
ServiceCenter
Management
Department
Exclusive
Director
Y.C Choi
And 3 others
Total 4person
Executive
Director
K.Y Hong
And 3 others
Total 4person
Vessel closure
Device
PCA
Disposable
Pressure
Transducer
Chief of
Department
M.S Choi
And 2 others
Total 3person
Service
Maintenance
Install
(All of us
Products)
Executive
Director
B.G Lee
And 8 others
Total 9person
Service
Maintenance
Upgrade
(GEN04,
GEN11,
Mega Power)
Chief of
Department
K.H Kim
Total 3person
Finance
Accounting
Purchasing
Inventory
Control
General Affairs
Executive
Director
H.W Lee
And 2 others
Total 3person
Patient Monitor
Anesthesia
Machines
OP Table
Surgical Light
Exclusive
Director
Y.C Choi
And 3 others
Total 4person
Plasma
Sterilizer(ASP)
Steam
Sterilizer
Washer
Notes: Management Department maintains quality function of the business. (This is listed from notes of 2019)
The organization chart is the same from 2019 except that marketing department director is now listed with associates.
No impact to Ethicon products .

## Slide 6
Organizational changes
Please answer the following: | Yes or No | Comments
Have there been any changes in the org chart since the last review? | Yes | One person has been reduced in the sales department and one person has been added in the marketing department.
Have there been changes in site location; expansion; or facility layout? | No. | N/A.
Have there been significant changes that adversely affect EES products, processes, or product release ? | No. | N/A.
Notes: The organization chart only list who works under who and how many so it possible this occurred but we cant see the associates by name only leadership.
This is how they always presented the data so I have no issue will look close at training.

## Slide 7
Follow-up Items from Pervious Review

## Slide 8: Follow –up Items from previous quality review
The follow-up to the annual business review was completed on 20-08-2019. The email at the time was attached below. | 
South Korea Team will update process/procedure to capture power supply serial numbers in repair records. | Responsible Party: Kim Kyung Ho Due date: 30-08-2019 Actions completed on (date): – South Korea Team supplied documentation for capturing serial numbers and power supplies. Completed 30-08-2019. Evidence is captured and will be documented in DOC023150 Rev C.
South Korea team will need to provide five more service records for review that includes: 1 Gen 04, 2 Gen 11 & 2 Megadyne Power 1000. | Responsible Party: Kim Kyung Ho Due date: 05-08-2019 Actions completed on (date): – South Korea Team supplied documentation for remaining service records for review. Completed 05-08-2019. Evidence is captured and will be documented in DOC023150 Rev C.
South Korea team will need updated records for: 1111225216_Repair_WO#285319 1111504010_Repair_WO#219132 for incomplete Yes/No box | Responsible Party: Kim Kyung Ho Due date: 05-08-2019 Actions completed on (date): – South Korea Team updated documentation for yes/no box of records for review. Completed 05-08-2019. Evidence is captured and will be documented in DOC023150 Rev C.
South Korea Team will need to update record: 1111143128_Repair _WO# 308124 for missing signature on the functional test for the supervisor. | Responsible Party: Kim Kyung Ho Due date: 05-08-2019 Actions completed on (date): – South Korea Team updated documentation for missing signature of record for review. Completed 05-08-2019. Evidence is captured and will be documented in DOC023150 Rev C.

## Slide 9: Follow –up Items from previous quality review
The follow-up to the annual business review was completed on 20-08-2019. The email at the time was attached below. | 
South Korea team will need updated management review slide for information related to service and repair activities: | Responsible Party: Kim Kyung Ho Due date: 05-08-2019 Actions completed on (date): – – South Korea Team supplied updated management review slide for activities related to service. Completed 05-08-2019. Evidence is captured and will be documented in DOC023150 Rev C.
Cincinnati Team need to provide SB# 18-003 for South Korea team to receive and confirm. | Responsible Party: James Swords Due date: 05-08-2019 Actions completed on (date): – WTCS Team supplied service bulletin to South Korea team confirmed thru e-mail. Completed 05-08-2019. Evidence is captured and will be documented in DOC023150 Rev C.
Cincinnati Team need to provide e-mail supplied by One MD to start capturing data in global service data base for all repairs. | Responsible Party: James Terry Due date: 01-08-2019 Actions completed on (date): –. James Terry supplied to South Korea team a e-mail about capturing data in MDS&R . Completed 01-08-2019. Evidence is captured and will be documented in DOC023150 Rev C.
Cincinnati team to follow up business management group to understand better there quality system requirements for internal audits. | Responsible Party:: Robert Peters Due date: 23-08-2019 Actions completed on (date): – Robert Peters supplied to South Korea team a e-mail about quality system requirements for internal audits . Completed 23-08-2019. Evidence is captured and will be documented in DOC023150 Rev C.

## Slide 10
Significant Changes in Quality System

## Slide 11: Significant Changes in Quality System
Please answer the following: | Yes or No | Comments
Has there been any changes significant changes to the Quality System in the last year? | No | N/A

## Slide 12
Average Monthly Service Volume During the Last 12 Months

## Slide 13: Average Monthly EES Service volume during last 10 months
Org Chart
Product Name | # of products serviced between 01-07-2019 to 30-04-2020
GEN11 | 778
GEN04 | 0
RF60 | N/A
MegaPower Generator | 33
MegaVac Smoke Evaluator | N/A
MegaSoft Pads | N/A
MiniVac | N/A
  81.8 units/month means Seolin Systems will be classified as a ‘Medium Volume’ Service Center

## Slide 14
Service Training Results

## Slide 15: Service Training Results for Newly Trained and Ongoing Personnel
Training Title/Type | Technician | Trainer (Person or System) | Training Date
GEN 11 Recertification | Kim,Kyungho Lee,Hyunjae Hwang,Daehwan | Hwang,Minho | 14-11-2019
Megadyne Recertification | Kim,Kyungho Lee,Hyunjae Hwang,Daehwan | Hwang,Minho | 14-11-2019
ISO Training | Kim,Kyungho Lee,Hyunjae Hwang,Daehwan Lee,Hanwoo | Hwang,Minho | 14-11-2019
Complaint Awareness | Kim.Kyungho Lee,Hyunjae Hwang,Daehwan Lee,Hanwoo | Summit | 28-10-2019
Notes: Trainer and trainee’s are the same. Hanwoo-Lee new for ISO and complaint, director of management department.
Question:
Is Gen 04 being serviced still re-certified and when, 0 volume from last review.
Answer: Yes Pm Only. Parts are not available to perform repair. See action slide.

## Slide 16: Service Training Results for Newly Trained and Ongoing Personnel
Product Name | Training the Trainer Name(s)
GEN11 | Kim,Kyungho
GEN04 | Kim,Kyungho
RF60 | N/A
MegaPower Generator | Kim,Kyungho Lee,Hyunjae Hwang,Daehwan
MegaVac Smoke Evaluator | N/A
MegaSoft Pads | N/A
MiniVac | N/A
Notes: Confirm these names.

## Slide 17
Product Quality Review

## Slide 18
Product Quality Review
Please answer the following: | Yes or No | Comments
Have the been any non-conformances for product or systems? | No | N/A.
Have there been any corrective/preventive actions identified in product or processes since the previous review? | No | N/A.
Has CAPA (Corrective And Preventive Actions) been determined as effective? | No | N/A.
Notes: No questions.

## Slide 19
Product Quality Review
    17 GEN11 & 3 Mega Power records were requested and reviewed, representing a sample of ANSI/ASQC Z1.4 Normal Level I AQL = 0.65 Accept 0; Reject 1 Inspection.
Notes: All Records completed by certified Repair tech from Korea Team.
No electrical safety performed due to being PM? Electrical Safety is not required for PM. No impact to product
This was a cleaning, no electrical or verification testing, only software upgrade? No output verification for software upgrades.
SN# 1111541148_ WO#381676 _PM_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. No issue. Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. No electrical safety on this since it was a PM. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause passement needs completed. 10/22/2019
SN# 1111637026_ WO#449326 _PM_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. No issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. No electrical safety on this since it was a PM. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause passement needs completed. 04/23/2020
SN# 1111729155_ WO#361525 _PM_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. No issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. No electrical safety on this since it was a PM. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause passement needs completed. 09/14/2019. Need to update asset for device on software.
SN# 1111540027_ WO#380485 _PM_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. No issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. No electrical safety on this since it was a PM. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause passement needs completed. 10/21/2019
SN# 1111519072_ WO#350425 _PM_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. No issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. No electrical safety on this since it was a PM. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause passement needs completed. 08/23/2019. Need to update asset for device on software.
SN# 1111508017_ WO#367786 _Complaint/ Not confirmed_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. No issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for PM. No electrical safety on this since it was a not confirmed. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause passement needs completed. No RA-is done with analysis and resolution codes. No-WOA- for decontamination and PM. No File is attached for customer sheet with both functional information and visual from the manual on their own form.10/10/2019.
SN# 1111234746_ WO#359955 _Complaint/ confirmed_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. Electrical safety performed values and specifications captured. No issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause passement needs completed. No-WOA- for decontamination. No File is attached for customer sheet with both functional information and visual from the manual on their own form or electrical.10/10/2019.
SN# 1111226326_ WO#429620 _ Complaint/ confirmed_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. Electrical safety performed values and specifications captured Replaced power supply and serial numbers captured. No issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The preliminary cause passement is completed. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. File attached is customer sheet with both functional information and visual from the manual on their own form and electrical safety. Part transfer complete with transactions. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time. No-WOA- for decontamination. No supervisor signature on attached electrical safety but it signed in scanned PDF.03/01/2020
SN # 1111226303_ WO#392375 _ PM_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. No issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. No electrical safety on this since it was a PM. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause assessment needs completed. 11/13/2019
SN # 1111225151_ WO#430780 _ PM_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. No issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. No electrical safety on this since it was a PM. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause assessment needs completed 02/22/2020
SN # 1111225085_ WO#406492 _ PM_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. No issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. No electrical safety on this since it was a PM. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause assessment needs completed 12/15/2019
SN # 1111218949_ WO/#413343 _ PM_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, /functional and QA form complete and final check list complete. No issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. No electrical safety on this since it was a PM. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause passement needs completed 01/08/2020
SN # 1111217635_ WO#424563 _ PM_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. No issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. No electrical safety on this since it was a PM. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause assessment needs completed 02/04/2020
SN # 111154007_ WO#349072 _ PM_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. No issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. No electrical safety on this since it was a PM. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause assessment needs completed 08/18/2019
SN # 1111148206_ WO#432273 _ compliant /confirmed Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. Electrical safety performed values and specifications captured Replaced main board and serial numbers captured. No issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The preliminary cause passement is completed. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. File attached is customer sheet with both functional information and visual from the manual on their own form and electrical safety. Part transfer complete with transactions. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time. No-WOA- for decontamination. No supervisor signature on attached electrical safety but it signed in scanned PDF. Why did you choose no failure found for preliminary cause assessment, a board was replaced? 03/19/2020
SN # 1111145865_ WO#450611 _ Complaint/ confirmed_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. Electrical safety performed values and specifications captured Replaced power supply and serial numbers captured. No issue. issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The preliminary cause passement is completed. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. File attached is customer sheet with both functional information and visual from the manual on their own form and electrical safety. Part transfer complete with transactions. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time. No-WOA- for decontamination. No supervisor signature on attached electrical safety but it signed in scanned PDF. Why did you choose no failure found for preliminary cause assessment, a board was replaced? 04/28/2020
SN # 1111130274_ WO#441812 _PM_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. Decontamination and receiving form completed, functional and QA form complete and final check list complete. issue Database work: The software is updated; the event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. No electrical safety on this since it was a PM. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause assessment needs completed 03/23/2020
SN # 175282001_ WO#365195_PM_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. functional form complete. Issue: No QA release but not required. No Decontamination and receiving form completed. Why? Done thru database. Why RGA not checked? RGA Needs to be N/A Database work: The event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. No electrical safety on this since it was a PM. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause assessment needs completed. The software needs captured. 09/19/2019
SN # 173884001_ WO#365195_Complaint/Confirmed_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. functional form complete. Issue: No QA release but not required. No Decontamination and receiving form completed. Why? Done thru database. Why RGA not checked? RGA Needs to be N/A Database work: The event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause assessment needs completed. The software needs captured. No part movement. Why ROHS leaded or ROHS not checked? This can occur on repair when you can physically look at the board.
SN # 152112001_ WO#429117_PM_ Certified tech: Hard copies GDP are good. Calibration and test device captured with description and due dates. functional form complete. Issue: No QA release but not required. No Decontamination and receiving form completed. Why? Done thru database. Why RGA not checked? RGA Needs to be N/A Database work: The event descriptions are in English. The conclusion code is done. The RA-is done with analysis and resolution codes. The IRO is complete but voided with no INV-for destination and source, so no way to confirm receiving other than paperwork provided. WOA- for decontamination and PM. File attached is customer sheet with both functional information and visual from the manual on their own form. No electrical safety on this since it was a PM. The actual hard copy they provided for this review would be a better attachment for the WO, it encompasses everything including test devices and receiving. Issue IRL need completed with INV and test devices need to be listed or paperwork attached, or they use ITH to save time would only need to attach electrical safety then. The preliminary cause assessment needs completed. The software needs captured. 02/16/2020

## Slide 20
Product Quality Review
Please answer the following: | Yes or No | Comments
Have there been any identified failures; breakdowns; or any outstanding issues related to spare part usage for this review? | No | N/A

## Slide 21
Product Quality Review
Service Bulletin # | Product | Service Bulletin Release Date | Implementation Date
SB20-003 | Service Manual V3 Release | 06-05-2020 | 08-05-2020
SB19-005 | Megapower Incorrect Authorized Rep on Rear Panel Label | 11-10-2019 | 30-01-2020
SB19-004 | Megadyne Field Action | 10-10-2019 | 30-01-2020
SB19-001 | Relase Megadyne CS-FRM-034 Rev 005 with Update Parts List | 15-08-2019 | 23-08-2019
Notes: Why so long on SB-19-005 was it us or them?
Answer: Extended time was needed to implement SB19-004 and SB19-005.

## Slide 22
Scorecard Metrics

## Slide 23
Scorecard Metric Review
Repair cycle time overall goal: 1 days. List repair cycle time overall goal above. | 
GEN11 | List results for GEN11 here. 1 days
GEN04 | List results for GEN04 here. N/A
RF60 | List results for RF60 here. N/A
MegaPower Generator | List results for MegaPower Generator here. 2 days
MegaVac Smoke Evaluator | List results for MagaVac Smoke Evaluator here. N/A
MegaSoft Pads | List results for the name of the train the trainer for MegaSoft Pads here. N/A
MiniVac | List results for MiniVac here. N/A
Metric | Results
Repeat service < 90 days, goal: ≤1 | Result = 1
Number of service complaints, goal: ≤1 | Result = 12
Customer Satisfaction goal: 2 Day | Result = 5 (ISO certified only)
Discuss timeliness for Global Scorecard reports. | Result = 100%
Discuss timeliness for Service Activity Reports | Result = 100%
Any additional metrics tracked (If service center tracks additional items) | N/A
Notes: Questions on complaints and customer stats results.
Answer: Implementation of new Megadyne products caused delay in cycle time due to the devices used in training while repairing. Service complaints were higher than expected due to the Korea team attempting to duplicate the customer’s error, unsuccessful.

## Slide 24
Management Review Summary

## Slide 25
Management Review Summary
List Results and Actions between 10-09-2018 to 30-07-2019 from Management Reviews: | 
Date of last management review | 30- 07-2019
Accomplishments? | # of hospitals to get GEN11 service: 178 hospitals (617 ea GEN11) # of service activity (CM: 112, PM 775, SW upgrade 41) Sales collaboration activity - One pager guide for failure: Analyze GEN11 cause of failure and provide guidance how to action to customer & sales team - Survey to know generator status by Hospital incl. competitors - 52 hospitals
Business plan goals? | Rapid service delivery to improve customer satisfaction Minimize customer inconvenience through regular preventive maintenance & rapid service within 1~2days Improve customer understanding by providing detailed service report Enhance equipment management thru regular customer training Expand KR Service Center role for sales team collaboration
Action items from the management review related to or affecting service? | N/A
Any significant Service Center internal audit trends? | N/A
Discuss any corrective and/or preventive actions implemented. | Was CAPA determined as effective? Yes___ No __N/A X
Other ? | N/A

## Slide 26
Compliance Update

## Slide 27
Compliance Update
Include Internal and Third Party:
  Any audits since last review – Yes X No___
    Last J&J audit took place on 07-11-2017 (Document in Korean)
  Any scheduled future audits – Yes X No___
    Future audit scheduled for 2020 date TBD
Notes: This will have some discussion, see notes from last year.
Per South Korea team the last audit was on 29-July-2017, for (Certification and Business Enhancement)
Next due date is 2020. Conducted every three years.
Not ISO Certified.
Question, With your Quality System in place how are yearly internal audits not being performed?
Per South Korea team the Management Department maintains quality function of the business and it requires audits once every three years.
See action item slide for follow up.
Last years response about internal Audits – HyeongJin, Associate Business Quality Manager, Business Quality Kim replied: Unfortunately, responsibility to perform internal audit by Seolin system is not addressed in either supplier quality agreement or any Korean regulation. However, Seolin system shall establish, maintain and improve the quality management system required for the support of products/services in accordance with the quality requirement of JJMK.

## Slide 28
Certifications

## Slide 29
Certifications
Please answer the following: | Yes or No | Comments
Are you certified by any Notified Body or Accreditation group? | No | N/A

## Slide 30
Records Review for Destruction

## Slide 31
Records Review for Destruction
  There are no records due for destruction at this time.
  GEN04 records dated 2014 will be eligible for destruction in 2029.

## Slide 32
General Discussion

## Slide 33
General
General information that additionally effects your business
  N/A

## Slide 34
Action Items
Below is a list of follow up action items as a result of the Seolin Systems South Korea Service Center 2020 Annual Business Review: | 
South Korea team to provide GEN04 Recertification documentation. | Responsible Party: South Korea team Due date: 03-06-2020 Actions completed on (02-06-2020): – E-mail sent with attachments from Min Ho containing the training records and test results. Evidence attached to DOC023150 Rev D
EES to provide information on how to identify the Megapower 1000 boards as leaded or ROHOS. (Only can be seen if board is viewed from opening the device.) | Responsible Party: James Terry Due date: 21-05-2020 Actions completed on (21-05-2020): – E-mail sent with attachment to Min Ho visual aid on how to identify the ROHO boards vs leaded. Evidence attached to DOC023150 Rev D.
Capturing serial number in service database for Mega Power 1000 boards. How to identify and locate serial number. | Responsible Party: James Terry Due date: 21-05-2020 Actions completed on (21-05-2020): – E-mail sent explaining to Min Ho that all boards purchased from the Cincinnati service center will have the serial number listed on the box and on the electrostatic bag that the board arrived in the QR code on the board can be scanned and will identify the number as well. Not all boards have QR codes. Those should have serial number listed again on box or bag. If the boards presently in South Korea were purchased from Megadyne and not the Cincinnati service center the QR code maybe the only way to identify the board. It more likely that this board asset is not captured in MDS& R. If there is no serial number on the box or bag this is not a spare part that is tracked in MDS&R. Evidence attached to DOC023150 Rev D

## Slide 35
Action Items – Continued
Below is a list of follow up action items as a result of the Seolin Systems South Korea Service Center 2020 Annual Business Review: | 
EES to supply sample of service file of receiving and Decontamination form Gen 11, So that can be supplied for the three Megapower records used for the review. SN # 152112001_ WO#429117, SN # 173884001_ WO#365195, 175282001_ WO#365195 | Responsible Party: James Terry Due date: 21-05-2020 Actions completed on (21-05-2020): – E-mail sent with attachment to Min Ho visual aid on what forms requested for Megapower unit. SN # 152112001_ WO#429117, SN # 173884001_ WO#365195, 175282001_ WO#365195. Evidence attached to DOC023150 Rev D
South Korea team to supply service file of Decontamination form for Megapower 1000. SN # 152112001_ WO#429117, SN # 173884001_ WO#365195, 175282001_ WO#365195 | Responsible Party: Min Ho Due date: 03-06-2020 Actions completed on (02-06-2020): – E-mail sent with attachments from Min Ho containing the three records showing decontamination. Evidence attached to DOC023150 Rev D
Determine why IRL are being voided in MDS&R for Inventory request orders. Determine if South Korea has INV# location to begin using MDS&R for tractability of product in and out of service center. | Responsible Party: Min Ho Due date: 30-06-2020 Actions completed on (): list what action was completed, by whom and list the objective evidence

## Slide 36
Action Items – Continued
Below is a list of follow up action items as a result of the Seolin Systems South Korea Service Center 2020 Annual Business Review: | 
Attach to MDS&R the following records for this quality review. SN# 1111508017_ WO#367786, SN# 1111234746_ WO#359955. Provided hard copies already. | Responsible Party: Min Ho Due date: 30-06-2020 Actions completed on (date): – list what action was completed, by whom and list the objective evidence
Re-open records in MDS&R the following records SN # 1111148206_ WO#432273, SN # 1111145865_ WO#450611 to update preliminary cause assessment to faulty parts for boards being replaced. | Responsible Party: Min Ho Due date: 30-06-2020 Actions completed on (date): – list what action was completed, by whom and list the objective evidence
Notes: Recommendations:
Try using Instruction test for capturing test data and test devices. This can be done by tools in MDS&R or in the WOA- Activity description field in the service database or attach actual output functional test form that already has test device, calibration due date and description listed. This is a requirement for all products and must be in the database.
Try to remember to update asset with current software for all products.
Remember if a failure is found on PM it must be converted to complaint, that includes updating interaction dates and contact with your tech and description fields locally.
On Megadyne forms N/A the RGA box if not applicable and N/A the ROHOS and leaded box.

## Slide 37
On-site Audit Required for Further Risk Assessment
EES to complete | Yes or No | Comments
Audit Required? | No | N/A.
Next quality review meeting 31-12-2021.