## Slide 1
Annual Quality Review Meeting (Emergent Medi-Tech (I) Pvt Ltd., India)
Date:04 – March - 2020
Ethicon Endo-Surgery

## Slide 2: Attendees
Name | Title
Mr. Ravindra Bhaiwal | Sr. Service Manager[J&J]
Mr. Suraj Kator | Asst. Manager [Emergent]
Mr. Nilesh Kanchan | Sr. Service Engineer [Emergent]
Shannon Gillespie | International Service Manager
Robert Peters | Service Quality Lead
James Swords | WTCS-Coordinator
James Terry | Product Release Technician
Agenda
Organizational changes
Follow-up items from previous quality review
Significant changes in Quality System
Average monthly service volume during last 12 months
Service training results
Product quality review
Scorecard Metrics
Management Review Summary
Compliance Update
Certifications
Records Review for destruction
General discussion

## Slide 3
Organizational Changes

## Slide 4
Organizational changes
Mr. Harish Rao
Director
Mr. Suraj Kator
Asst. Manager
Mr. Nilesh Kanchan
Sr. Service Engineer
Mrs. Aruna Aagare
Administrator / MR
Mr. Akshay Hejmadi
Call Centre Representative
Mr. Tanmay Gaonkar Admin-eWay bills and complaint login
Mr. Omprakash Kori
Service Centre Assistant
Mr. Ravindra Bhaiwal
Sr. Service Manager (J&J has oversight of Emergent.)
Notes: Titles have changed for the following associates from 2019: Suraj Kator was Sr. Customer support engineer. Nilesh Kanchan was customer support Engineer.
No longer listed is Mr. Viren Parelkar, he was the general manager. See next slide for new associates.

## Slide 5
Organizational changes
Please answer the following: | Yes or No | Comments
Have there been any changes in the org chart since the last review? | Yes | Mr. Harish Rao is the company director, taking charge of service center w.e.f. October 2019 . Mrs. Aruna Aagare was appointed as Administrator-Service Center &MR w.e.f.April 2019 Mr. Akshay Hejmadi was appointed as Call Center Executive w.e.f. April 2019. Mr. Tanmay Gaonkar is appointed for complaint handling executive w.e.f. Feb.2020
Have there been changes in site location; expansion; or facility layout? | No | N/A.
Have there been significant changes that adversely affect EES products, processes, or product release ? | No | N/A
Notes: Question: What about Mr. Omprakash Kori. Does this associate have any duties of repair for Ethicon products?
Where is the organization chart for all associates related to Hong Kong, Thailand, Philippines, Vietnam, Malaysia See next two slides from 2019
Answer: Mr Omprakash Kori does not have any duties for repair of Ethicon Capital
organization chart for all associates related to Hong Kong, Thailand, Philippines, Vietnam, Malaysia See next two slides

## Slide 6
Organizational changes
Country | Engineers | Certification
India | Nilesh Kanchan, Suraj Kator | GEN11,GEN04,RF60
Hong Kong | William So | GEN11PM,GEN04PM,RF60PM
Thailand | Chawalit Orraeiam, Chonlathit Traewachiwiphak | GEN11PM
Philippines | Honorio Castro, Elmer Ordono | 
Vietnam | Duc Min Nguyen, Dao Xuan Tu, Dang Van Phuong, Dao Cong Thanh | 
Brunei Darussalam | Alex Chan Swee Yen | 
Malaysia | Dhurai Raj A. Selvarajoo, Wang Hooi Peng | 

## Slide 7
Follow-up Items from Previous Review

## Slide 8: Follow –up Items from previous quality review
Below is a list of follow up action items as a result of the (Emergent Medi-Tech (I) Pvt. Ltd. India) 2019 Annual Business Review: | 
List action item | None

## Slide 9
Significant Changes in Quality System

## Slide 10: Significant Changes in Quality System
Please answer the following: | Yes or No | Comments
Has there been any changes significant changes to the Quality System in the last year? | No | N/A

## Slide 11
Average Monthly Service Volume During the Last 12 Months

## Slide 12: Average Monthly EES Service volume during last 12 months
Org Chart
Product Name | # of products serviced between 01-March-2019 TO 12-Feb-2020 |  | 
 | INDIA | APAC | TOTAL
GEN11 | 167 | 372 | 539
GEN04 | 04 | 00 | 04
RF60 | 00 | 00 | 00
MegaPower Generator | 34 | 00 | 34
MegaVac Smoke Evaluator | 00 | 00 | 00
MegaSoft Pads | 00 | 00 | 00
MiniVac | 00 | 00 | 00
577 Total units represents an average of 48.08 units/month (1st March 2019 – 1st Feb 2020). This data represents 12 months from last review. This means Emergent will remain a Medium-volume Service Center.
Notes: From 2019 Review was the Gen 04 Only PM. Since in 2019 no parts were available.
Answer: The GEN04’s were not PM’s, they were repairs for parts that Emergent had remaining in stock. No main PC boards are currently available.
Do these totals include the field service work from Hong Kong, Thailand, Philippines, Vietnam, and Malaysia?
Answer: Slide has been updated with applicable countries volumes for PM’s.

## Slide 13
Service Training Results

## Slide 14: Service Training Results for Newly Trained and Ongoing Personnel
Document New/Ongoing/Re-Certification training in chart for each product here.
Note: Any trained employee should have training records available upon request for each product serviced.
Training Title/Type | Technician |  | Trainer (Person or System) | Training Date
GEN 11 Recertification | Suraj Kator Nilesh Kanchan William So [PM] Chawalit Orraeiam [PM] Duc Min Nguyen [PM] Elmer Ordono [PM] Dhurai Raj [PM] | Dao Cong Thanh [PM] Dang Van Phuong [PM] Dao Xuan Tu [PM] Alex Chan Swee Yen [PM] Chonlathit Traewachiwiphak [PM] Larissa Loh Jia Min [PM] Wang Hooi Peng [PM] | Mr. Harish Rao | (01-JAN-2020)
Gen 04 Recertification | Suraj Kator Nilesh Kanchan William So [PM] |  | Mr. Harish Rao | (01-JAN-2020)
ISO Training | N/A |  | N/A | N/A
Complaint Awareness | Suraj Kator Nilesh Kanchan William So Chawalit Orraeiam Duc Min Nguyen Elmer Ordono Dhurai Raj | Dao Cong Thanh Dang Van Phuong Dao Xuan Tu Alex Chan Swee Yen Chonlathit Traewachiwiphak Larissa Loh Jia Min Wang Hooi Peng | Mr. Harish Rao | (01-JAN-2020)
Notes: No longer listed Honorio Castro.
Answer: Gentlemen is no longer with the company.
Question GEN 04 no longer listed was GEN04 training completed prior to last review and not included?
Answer: Updated within slide.
Question was Complaint awareness performed in Summit for 2020? Does Emergent and the other countries have access to Summit for training?
Answer: Yes.

## Slide 15: Service Training Results for Newly Trained and Ongoing Personnel
Product Name | Training the Trainer Name(s)
GEN11 | Suraj Kator
GEN04 | Nilesh Kanchan
MegaPower Generator | Suraj Kator
MegaVac Smoke Evaluator | Nilesh Kanchan
MegaSoft Pads | No Training Executed
MiniVac | No Training Executed
Notes: Mr. Suraj Kator is the trainee trained by EES Associate Bill Mead.

## Slide 16
Product Quality Review

## Slide 17
Product Quality Review
Please answer the following: | Yes or No | Comments
Have there been any non-conformances for product or systems? | Yes. | During Annual Calibration procedure in the month of Nov.2019, HyPot Tester, Torque Drivers, Digital Oscilloscope went out of acceptable range.
Have there been any corrective/preventive actions identified in product or processes since the previous review? | Yes | Service Center has already procured new HyPot tester and Torque Drivers & Digital Oscilloscope is sent for repairs.
Has CAPA (Corrective And Preventive Actions) been determined as effective? | Yes | We had reported the NC as soon as the event happened, yet to report the Corrective Action to validate the same. As a preventive action,during the year 2020, Service Center will have Standby testing Equipment to overcome any such situations and ensure un-interrupted functioning and Service to Customers.
Notes: Question: What is the status of the corrective action? Quality has seen no communication addressing the effect of product that was accepted on the out of calibration equipment. Specifically the hypot tester.
Answer: See action item slide for evidence.

## Slide 18
Product Quality Review
A set of twenty service records were selected at random and reviewed based on ANSI/ASQC Z1.4 Normal Level I AQL = .65 - Accept 0; Reject 1
  12 GEN11 Records
  7 Mega Power Records
  1 Gen 04
Notes: Summary:
After looking over the files sent as hard copies. GDP are very well done. Frm003176 was completed and correlated to the hard copies.
What was done well: All documentation complete including electrical safety calibration devices and due dates Including QA data (if applicable). The database was completed very well for all Gen 11 and Megadyne as well as Gen 04 for RA- subject codes and analysis codes addressed. All test data was attached. IRL completed for Gen 11 and some Megadyne. They also pick their test equipment in the database.
Need to discuss: The database was system requirements and IRO mostly and IRL for Megadyne and Gen 04. Need to discuss ITH related to Gen 11 and Megadyne (1 record), using ours that reference PR000920 and PR001566.
Recommendations: File name attached contain work order number. Per 100549921 and 100549929, we need to talk to them about create WOA-CM/PM and associate the RA-work under that WOA. Talk about updating assets and issue with overriding account names.
Result: India team will complete IRO and IRL for Megadyne products. Discussed CM creation as new requirements are coming in harmonization procedures for service and repair. Discussed issue of asset updating.
Sn# 111139505_ EN/0319/002 WO#261791_ Complaint_ Confirmed. _Repair completed by approved Technician. Power supply replaced. Documented Serial Number.FRM003176_Complete. For Repair all documentation complete including electrical safety, calibration devices and due dates and technical bulletins and QA. GDP On documents look good. IRO completed. Line completed and shipped. Ra Completed and subject codes and analysis codes addressed. Issue: Used an ITH that referenced our PR000920. Just pick global one. File Name of attachment.
Sn# 111138890_ EN/0619/021 WO#292142_ Complaint_ Confirmed. _Repair completed by approved Technician. Power supply replaced. Documented Serial Number.FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates and technical bulletins and QA. GDP On documents look good. IRO completed. Line completed and shipped. Ra Completed and subject codes and analysis codes addressed. Issue: Used an ITH that referenced our PR000920. Just pick global one. File Name of attachment.
Sn# 1111410222_ EN/0619/009 WO#318497_ Complaint_ Confirmed. _Repair completed by approved Technician. Main board replaced. Documented Serial Number.FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates and technical bulletins and QA. GDP On documents look good. IRO completed. Line completed and shipped. Ra Completed and subject codes and analysis codes addressed. Issue: Used an ITH that referenced our PR000920. Just pick global one. File Name of attachment.
Sn# 1111834142_ EN/0819/020 WO#322219_ Complaint_ Confirmed. _Repair completed by approved Technician. Cooling fan error. FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates and technical bulletins and QA. GDP On documents look good. IRO completed. Line completed and shipped. Ra Completed and subject codes and analysis codes addressed. Issue: Used an ITH that referenced our PR000920. Just pick global one. File Name of attachment.
Sn# 1111511047_ EN/0719/007 WO#325664_ Complaint_ Confirmed. _Repair completed by approved Technician. LCD Display replaced. FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates and technical bulletins and QA. GDP On documents look good. IRO completed. Line completed and shipped. Ra Completed and subject codes and analysis codes addressed. Issue: Used an ITH that referenced our PR000920. Just pick global one. File Name of attachment.
Sn# 1111832046_ EN/0919/007 WO#338249_ Complaint_ Confirmed. _Repair completed by approved Technician. LCD Display replaced. FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates and technical bulletins and QA. GDP On documents look good. IRO completed. Line completed and shipped. Ra Completed and subject codes and analysis codes addressed. Issue: Used an ITH that referenced our PR000920. Just pick global one. File Name of attachment.
Sn# 1111229885_ EN/0919/012 WO#355029_ Complaint_ Confirmed. _Repair completed by approved Technician. Main board replaced. Documented Serial Number .FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates and technical bulletins and QA. GDP On documents look good. IRO completed. Line completed and shipped. Ra Completed and subject codes and analysis codes addressed. Issue: Used an ITH that referenced our PR000920. Just pick global one. File Name of attachment.
Sn# 1111702067_ EN/0919/020 WO#355619_ Complaint_ Confirmed. _Repair completed by approved Technician. Cable assembly receptacle .FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates and technical bulletins and QA. GDP On documents look good. IRO completed. Line completed and shipped. Ra Completed and subject codes and analysis codes addressed. Issue: Used an ITH that referenced our PR000920. Just pick global one. File Name of attachment.
Sn# 1111606134_ EN/0919/033 WO#365549_ Complaint_ Confirmed. _Repair completed by approved Technician. LCD Display replaced.FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates and technical bulletins and QA. GDP On documents look good. IRO completed. Line completed and shipped. Ra Completed and subject codes and analysis codes addressed. Issue: Used an ITH that referenced our PR000920. Just pick global one. File Name of attachment.
Sn# 1111408404_ EN/0419/013 WO#276655_ Complaint_ Confirmed. _Repair completed by approved Technician. Main board replaced.FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates and technical bulletins and QA. GDP On documents look good. IRO completed. Line completed and shipped. Ra Completed and subject codes and analysis codes addressed. Issue: Used an ITH that referenced our PR000920. Just pick global one. File Name of attachment.
Sn# 1111437090_ EN/0919/029 WO#361768_ Complaint_ Confirmed. _Repair completed by approved Technician. Main board replaced and power supply.FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates and technical bulletins and QA. GDP On documents look good. IRO completed. Line completed and shipped. Ra Completed and subject codes and analysis codes addressed. Issue: Used an ITH that referenced our PR000920. Just pick global one. File Name of attachment.
Sn# 1111318058_ EN/0519/009 WO#296374_ Complaint_ Confirmed. _Repair completed by approved Technician. Main board replaced .FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates and technical bulletins and QA. GDP On documents look good. IRO completed. Line completed and shipped. Ra Completed and subject codes and analysis codes addressed. Issue: Used an ITH that referenced our PR000920. Just pick global one. File Name of attachment.
Sn# GN4111159_ EN/0519/009 WO#349814_ Complaint_ Confirmed. _Repair completed by approved Technician. Bezel assembly .FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates. GDP On documents look good. Completed and subject codes and analysis codes addressed. Issue: No IRO created. No IRL created or shipped. File Name of attachment.
Sn# 180089001_ EN/0819/002 WO#325498_ Complaint_ Confirmed. _Repair completed by approved Technician. Main board replaced .FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates. GDP On documents look good. Ra Completed and subject codes and analysis codes addressed. Issue: Used an ITH that referenced our PR001566. Just pick global one. No IRO created. No IRL created or shipped.
Sn# 175330001_ EN/0519/001 WO#339571_ Complaint_ Confirmed. _Repair completed by approved Technician. Main board replaced .FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates. GDP On documents look good. Ra Completed and subject codes and analysis codes addressed. Issue:. Just pick global one. No IRO created. No IRL created or shipped.
Sn# 175340001_ EN/119/022 WO#330051_ Complaint_ Confirmed. _Repair completed by approved Technician. Main board replaced .FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates. GDP On documents look good. Ra Completed and subject codes and analysis codes addressed IRO completed. Line completed and shipped. Issue: Just pick global one.
Sn# 181271001_ EN/0120/012 WO#415557_ Complaint_ Confirmed. _Repair completed by approved Technician. Front panel PCB board .FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates. GDP On documents look good. Ra Completed and subject codes and analysis codes addressed. Issue: Just pick global one. No IRO created. No IRL created or shipped.
Sn# 182092001_ EN/1019/031 WO#378399_ Complaint_ Confirmed. _Repair completed by approved Technician. Front panel PCB board .FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates. GDP On documents look good. Ra Completed and subject codes and analysis codes addressed. Issue: Just pick global one. No IRO created. No IRL created or shipped.
Sn# 180108801_ EN/1019/024 WO#347231_ Complaint_ Confirmed. _Repair completed by approved Technician. Main board .FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates. GDP On documents look good. Ra Completed and subject codes and analysis codes addressed. Issue: Just pick global one. No IRO created. No IRL created or shipped.
Sn# 182111001_ EN/1019/020 WO#368878_ Complaint_ Confirmed. _Repair completed by approved Technician. Power supply .FRM003176_Complete. For Repair all documentation complete including electrical safety calibration devices and due dates. GDP On documents look good. Ra Completed and subject codes and analysis codes addressed. Issue: Just pick global one. No IRO created. No IRL created or shipped.

## Slide 19
Product Quality Review
Please answer the following: | Yes or No | Comments
Have there been any identified failures; breakdowns; or any outstanding issues related to spare part usage for this review? | Yes | 3 No of Main PCB Assembly of GEN11 were found to be faulty on arrival (Part No.EE01-63095)

## Slide 20
Product Quality Review
Service Bulletin # | Product | Service Bulletin Release Date | Implementation Date
SB19-001 | Megadyne | 29-Jul-2019 | 22-Aug-2019
SB19-004 | Megadyne | 26-Sep-2019 | 14-Oct-2019
SB19-005 | Megadyne | 17-Jan-2020 | 30-Jan-2020
Notes: Confirmed,

## Slide 21
Scorecard Metrics

## Slide 22
Scorecard Metric Review
Repair cycle time overall goal: 10 days List repair cycle time overall goal above. | 
GEN11 | 22 days
GEN04 | 1 Day
MegaPower Generator | 24 days
MegaVac Smoke Evaluator | NA
MegaSoft Pads | NA
MiniVac | NA
Metric | Results
Repeat service < 90 days, goal: ≤1 | Result = 2
Number of service complaints, goal: ≤1 | Result = 0
Customer Satisfaction goal: 90%_ | Result = 80%
Discuss timeliness for Global Scorecard reports. | Result = 100%
Discuss timeliness for Service Activity Reports | Result = 100%
Any additional metrics tracked (If service center tracks additional items) | N/A
Notes: Question: Can you please explain why the goals were not met for devices and customer sat?
Answer: Per team, awaiting for customer quote approval and PO and Work Order re-open process. The reason for an increased cycle time on the GEN11 was due to tool calibration. The Mega Power 1000 increase was due to lack of Spare parts. This also impacted customer sat.

## Slide 23
Management Review Summary

## Slide 24
Management Review Summary
List Results and Actions between 18-09-2018 to 20-09-2019 from Management Reviews: | 
Date of last management review | 20-09-2019
Accomplishments? | Not related to Service Center
Business plan goals? | Not related to Service Center
Action items from the management review related to or affecting service? | Not related to Service Center
Any significant Service Center internal audit trends? | No internal audit trends detected
Discuss any corrective and/or preventive actions implemented. | N/A
Other ? | N/A

## Slide 25
Compliance Update

## Slide 26
Compliance Update
Include Internal and Third Party:
  Any audits since last review – Yes_X_ No___
      Successful J&J compliance audit was conducted on 26th, Sep., 2019 by Mr. Bhuwandeep Singla. There were 6 findings related to service and repair. All findings were closed.
      Internal audit was conducted on 23/Apr/2019. No findings related to service center.
  Any scheduled future audits – Yes_X_ No___
      Internal audit will be performed on 22/Apr/2020.
Notes: Question: Can we know more about the six findings? Please specify what the topic was for each of the 6 findings and whether they were major or minor.
Answer: See action slide and evidence.

## Slide 27
Certifications

## Slide 28
Certifications
(ISO Certification 9001:2015 | Certificate No.: IN86136A
dated 26/July/2019 valid up to 25/July/2022)
Please answer the following: | Yes or No | Comments
Are you certified by any Notified Body or Accreditation group? | Yes | ISO certification 9001:2015
Notes: Confirmed.

## Slide 29
Records Review for Destruction

## Slide 30
Records Review for Destruction
No Records for destruction at this time. Pegasus no applicable to this service site.

## Slide 31
General Discussion

## Slide 32
General
No additional discussion was needed.

## Slide 33
Action Items
Below is a list of follow up action items as a result of the (Emergent Medi-Tech (I) Pvt. Ltd. India) 2020 Annual Business Review: | 
Supply information on finding for Hypot tester | Responsible Party: Suraj Kator Due date: 19-03-2020 Actions completed on (date): – Suraj Kator supplied documentation and explanation for the finding for the hypot tester. Completed 06-03-2020 Evidence is captured and will be documented in supporting documents folder.
Supply information for the 6 findings from audit observation | Responsible Party: Suraj Kator Due date: 19-03-2020 Actions completed on (date): – Suraj Kator supplied documentation and explanation for the finding for the audit observation. Completed 06-03-2020 Evidence is captured and will be documented in supporting documents folder.
Supply information on the service volumes for the APAC countries | Responsible Party: Suraj Kator Due date: 19-03-2020 Actions completed on (date): – Suraj Kator updated slide and embedded in this presentation. Completed 06-03-2020 – See slide 13 of this presentation.
Supply information on the recertification for GEN04 and PM’s | Responsible Party: Suraj Kator Due date: 19-03-2020 Actions completed on (date): – Suraj Kator updated slide and embedded in this presentation. Completed 06-03-2020. – See slide 15 of this presentation.
Notes: Recommendation: James Terry to assist in supplying Global instruction test for GEN11 and Megapower 1000.
Evidence is captured and will be documented in

## Slide 34
On-site Audit Required for Further Risk Assessment
EES to complete | Yes or No | Comments
Audit Required? | No | N/A.
Next quality review meeting 31-December 2021.