## 3.1 Specifications
Table 1 Specifications and basic dimensions of trocars (Units: mm) |  |  |  |  | Please Click the following link for detailed dimension analysis per drawing specification
Specification code (list) | Inner diameter of puncture set cannula |  | Working length of puncture set |  | Dimension Analysis for XCEL Trocars
| Nominal value | Limit deviation | Nominal value | Tolerance
3 | Φ≤3 | ＋0.32 | 50≤L≤200 | ±1% | Dimension Analysis for Optiview Trocars
0
5 | 3＜Φ≤6 | ＋0.34 |  |  | Dimension Analysis for BASX Trocars
0
7 | 6＜Φ≤10 | ＋0.36
0
12 | 10＜Φ≤14 | ＋0.38
0
15 | 14＜Φ≤18 | ＋0.40
0
Note: The specification codes of trocars are not limited to those listed here. Other specification codes that are not listed here but meet the requirements of this standard may also be used.
Specification code 3 is not applicable as ethicon products are offered from 5 mm and above . - Meets the tolerance requirement as per YY standard for Xcel, Optiview and BASX trocars. - Tolerance Range is 0.05mm to 0.08mm for Inner diameter of puncture set cannula against the YY requiremment of 0.32mm to 0.4mm - Tolerange range is 0.25mm to 0.38mm against the YY requirement of 1% for Working length of puncture set Please refer attached files for dimensional verification
| OD and ID dimensions of sleeve and OD of Obturator were compared with GB/T 14486-2008.(For effective length theoretical analysis please refer the links above in Column G.) |  |  |  | Attachment-08

## 3.2 Material
3.2 Materials See Table 2 for the manufacturing materials for the major components of trocar.
Trocar component | Material designation | Standard code | Drawing No. | R-Spec No. | Material | Matches Standard Code | Remarks
Obturator | PC | YY/T 0806 | D12553 Rev U | R1002-58 R1002-13 R1002-32 | Polycarbonate | Yes
| ABS | GB/T 12672 | N/A
| 12Cr18Ni9, 06Cr19Ni10, 20Cr13 | GB/T 1220 | D12551 Rev J | R2004-01 Cond F | UNS S30400 | Yes | Chemical Composition is slightly different.
Cannula | PC | YY/T 0806 | D28771 Rev H | R1002-13 | Polycarbonate | Yes
Paracentesis knife | 12Cr18Ni9, 06Cr19Ni10, 20Cr13 | GB/T 3280 | D02138 Rev L | R2001-01 Cond A R2002-02 Cond F | UNS S44002 SUS 420J2 | Yes | Chemical Composition is slightly different.
Note: Other materials that are evaluated as safe and in line with the requirements of this standard can also be used.

## 4 Requirement - 5 Test Methods
Clause No. | Draft YY Requirement/Testing | XCEL Supporting Evidence | XCEL-Optiview Supporting Evidence | BASX Supporting Evidence | Match | Justification/Remarks | Update | Attachments
4.1 | Appearance
4.1.1 | The external surface of the trocar shall be smooth and flat without burr, and shall be free from such defects as bubbles, impurities, crack, hole, groove and sinters. | Specification: MS00131-Rev X, SPE000171-Rev K-Section 4.1.1, 4.1.2, 4.1.3 & 4.1.4 | Specifications: Material Specification: 5mm Trocars: SPE002054 Rev C 12 mm Dilating Tip: SPE002433 Rev B, 12 mm Bladeless & Blunt tip: SPE002432 Rev B, Clause 4.1.1,4.1.2,4.1.3 & 4.1.4 | Specifications: SPE000623 Rev G (4.1.1, 4.1.2, 4.1.3) | Yes | The Material Specification doument captures the identification, categorization and allowances for such defects as mentioned in the draft YY Standard. Devices are inspected for these defects during the Finished Goods Quality Assurance inspections.
4.1.2 | Accumulation of lubricant visible to the naked eye is not allowed on the surface of trocar or in the simulated use of trocar. | Specification: MS00131-Rev X, SPE000171-Rev K-Section 4.1.1, 4.1.2, 4.1.3 & 4.1.4 | Specifications: Material Specification: 5mm Trocars: SPE002054 Rev C 12 mm Dilating Tip: SPE002433 Rev B, 12 mm Bladeless & Blunt tip: SPE002432 Rev B, Clause 4.1.1,4.1.2,4.1.3 & 4.1.4 | Specifications: SPE000623 Rev G (4.1.1) | Yes | The material specification doument captures the identification, categorization and allowances for such defects as mentioned in the draft YY Standard. Devices are inspected for these defects during the Finished Goods Quality Assurance inspections.
4.1.3 | The specification mark on the trocar body shall be clearly visible. | Specification: MS00131-Rev X, SPE000171-Rev K-Section 4.2 | Specifications: Material Specification: 5mm Trocars: SPE002054 Rev C 12 mm Dilating Tip: SPE002433 Rev B, 12 mm Bladeless & Blunt tip: SPE002432 Rev B, Clause 4.2 | Specifications: SPE000623 Rev G (4.2) | Yes | The material specification doument captures the identification, categorization and allowances for such defects as mentioned in the draft YY Standard. Devices are inspected for these defects during the Finished Goods Quality Assurance inspections.
4.1.4 | If a paracentesis knife is attached at the head end of the obturator, the paracentesis knife shall be flat and free of rust, sharp edge, burr and obvious pitting; the cutting edge of paracentesis knife shall be free of gap, white cast, wire edge, crack, etc. | Specification: See Note 5 of 5mm Blade- D02138P03-Rev L(D5LT, D5ST, 2D5LT,2D5ST). 11mm Blade- D02327P02-Rev D (D11LT, K11LT) 12mm Blade- D02837P02-RevC: Blade Drawing for 512SD,512XD, D12LT, D12XT, K12LT.2D12LT,2D12XT R-Spec-R2002-02 Cond F Component Approval Document: CA009742, CA058891 |  | BASX Trocar do not have paracentesis knife, thus this requirement is not applicable for BASX Trocars CTB5LT/CTB5ST: D17821 Rev F CTB11LT: D17822 Rev G CTB12LT: D30988 Rev B TB5LT/TB5ST: D16973 Rev H TB11LT:D17325 Rev J TB12LT/TB12ST: D17057 Rev K | Yes | Such component level defects are captured in the engineering drawing requirements and are inspected during the manufacturing processes at supplier end. Inspection reports are submitted to validate that the engineering drawing requirements were met.
5.1 | Using normal or corrected visual acuity, make visual observation under illuminance of 300 Lx~750 Lx and hand touching for examination; the result shall comply with the provisions of 4.1. | Test method: I8533 Rev AD, PR000316- Rev AB section 4.3 | Test method: Visual Inspection 5 mm Trocar PR001213 Rev B Section 4.1 12 mm Trocar PR001122 Rev C Section 4.1 12 mm Trocar (2D12LT/2D12XT) PR001124 Rev F Section 4.1 | Visual Inspection PR000471 Rev G | Yes | Tests related to Visual Inspection are done on products on completion of the entire assembly process, in well lit rooms with ample amount of light to visually inspect the components for defects mentioned in the requirement. The general measure of illuminance for a well lit room or laboratory is approximately 500 lux. | The specific value of illumination can be suggested to be removed from the test method clause as the illumination index of the lab in which the tests performed are not officially recorded in the test reports. | Attachment 12 (Data collected from internet)
4.2 | Dimensions
| The cannula inner diameter Φ of trocar and the working length L of the trocar shall comply with the provisions in Table 1. | Refer Tab "3.1 Specifications" for details
5.2 | Use a general/special measuring tool to make measurement; the results shall comply with the provisions of 4.2. | Test method followed matches that of draft YY Standard | Test method followed matches that of draft YY Standard | Test method followed matches that of draft YY Standard | Yes | Dimensions are measured using standard measurement tools
4.3 | Surface Roughness
| If a paracentesis knife is attached at the head end of the obturator, the surface roughness Ra of paracentesis knife shall be ≤ 0.8 μ m. | Specification: See Note 5 of 5mm Blade- D02138P03-Rev L(D5LT, D5ST, 2D5LT,2D5ST). 11mm Blade- D02327P02-Rev D (D11LT, K11LT) 12mm Blade- D02837P02-RevC: Blade Drawing for 512SD,512XD, D12LT, D12XT, K12LT.2D12LT,2D12XT Note Item: 3 - Surface Finish Max 63 µinch (1.6µm) Test Report: Attachment 1 FAIs are present in the respective sharepoint Evidence Folder |  | BASX Trocar do not have paracentesis knife, thus this requirement is not applicable for BASX Trocars | Yes | Supplier measurement reports shows the surface measurement is below <0.8µm and meets the surface roughness as per YY requirement. Polish Side #1 (Max Value): 0.093µm Polish Side #2 (Max Value): 0.090µm Tip Edge (Max Value): 0.688µm (Refer Attachment 1 for details) Engineering Studies will be prepared to validate the capability of the manufacturer to maintain the surface roughness value within the limit mentioned in draft YY standard, since Ethicon Specification for surface roughness in Engineering drawings is 1.6µm |  | Attachment 1
5.3 | The surface roughness shall be inspected by adopting a comparison specimen; the result shall comply with the provisions of 4.3. | Test method followed matches that of draft YY Standard | Test method followed matches that of draft YY Standard | BASX Trocar do not have paracentesis knife, thus this requirement is not applicable for BASX Trocars | Yes | Surface roughness measurement is done with help of surface roughness tester which is a standard industry method. |  | Attachment 17 (Test Method followed by Supplier)
4.4 | Flexibility
4.4.1 | The gas injection valve of the trocar shall be flexibly opened/closed, without blocking or jamming. | Specifications: DR000111-Rev E- 039, DS0110-Rev K-051, DS0100-Rev J-050, DR000083-Rev H-038 Reports: DV Test Report- EP3-0474 Rev B, EP3-0475 Rev B, FAI Stop Cock Assembly(D13142): CA055881 Rev A | Specifications: 5 mm Trocar DR00239 Rev J Sub Assembly requirement ID 0336 12 mm Trocar DR00249 Rev E- Sub Assembly requirement ID SALR-08.2 Reports DV Test Report- EP3-0474 Rev B, EP3-0475 Rev B, FAI Stop Cock Assembly(D13142): CA055881 Rev A | Specifications: DR000164 Rev E-Performance Requirement ID 2-57 Reports: Functionality Testing: PSP002432 Rev A Page 55 of 71 Functionality Testing Memo: SCN004383 Rev A DV Testing Protocol: PRC014322 Rev A FAI Stop Cock Assembly (D13142 Rev K): CA006513 Rev A | Yes | As per design specifications, the requirement is that the Force To Actuate the insufflation valve shall not exceed 2.5lbf-in (0.29 N-m approx.) when rotated in both a clockwise and anti-clockwise direction The requirement captures the required flexibility of the mechanism, and quantifiable values which are set based on voice of customer, market surveillance and research makes the results stricter and more reliable in meeting the standards.
4.4.2 | If the sheath cap is detachable, its assembly and disassembly shall be flexible and convenient and shall be free from phenomenon of obstruction or blocking. | Specification: DR000111 – Rev E- Req. ID 37, 116 DS0110- Rev K-Req. Id 032, 050 & DR000083- Rev H-Req. ID 025, 035 Reports: PRC001280A- 11mm and 12mm Outer Seal Related Design Verification Testing | Specification: 12 mm Trocar: DR000249 Rev E System Requirement ID SALR-07.3, SLR-08.5 Reports PRC058935 PRC053519 | There is no sheath cap in BASX Trocar, hence this clause is not applicable CTB5LT/CTB5ST: D17821 Rev F CTB11LT: D17822 Rev G CTB12LT: D30988 Rev B TB5LT/TB5ST: D16973 Rev H TB11LT:D17325 Rev J TB12LT/TB12ST: D17057 Rev K | Yes | Source specifications describe the requirement as: the force required to actuate the latching ring to attach or remove the Universal Cap from the sleeve housing shall be greater than 0.36 lbf and less than 6.84 lbf. with reliability of 99% at 95% confidence interval. Integrated Removable Seal (Universal Reducer Cap) To Trocar Housing Attachment: The force to remove the Integrated Removable Seal (Universal Reducer Cap) from the sleeve housing of an ENDOPATH® Xcel 12 mm surgical trocar must be greater than 17.5 lbf. when loading is applied axial to the Trocar and the Integrated Removable Seal (Universal Reducer Cap). Rationale: The tensile axial force requirement is a function of the maximum trocar retention force for the Endopath XCEL Bladeless Trocar with stability sleeve. The requirement ensures that the assemby and disassembly of the sheath cap (Universal Seal Cap) is flexible and free of any blockage. The requirement values are set as per voice of customer, market surveillance and research. Thus, tests done to quantify the requirement are stricter and also more reliable.
5.4 | Simulate the usage action and make inspection through hand feeling; the result shall comply with the provisions of 4.4. | Test Method: Load Cells SCN019439B – Test method for force to install and removal of Integrated Removable Seal SCN019444 TR11513 | Test Method: Load Cells SCN019439B – Test method for force to install and removal of Integrated Removable Seal SCN019444 TR11513 | There is no sheath cap in BASX Trocar, hence this clause is not applicable | No match but better and quantifiable test method | Source specifications testing is done using load cell arrangement, whereas draft standard mentions simulation of actual operation. The source specification testing is quantifiable, stricter and thus more reliable.
4.5 | Matching Performance
4.5.1 | The puncture set shall be well fit with the obturator without any blocking when being inserted and extracted. | Specifications: DS0100-Rev J-Req. ID 008, DS0110-Rev K-Req. ID 008 DR000204-Rev F- Req. ID SYS-008 Reports: QA0081-Rev F –Page 11 of 44 QA0085-Rev L – Page 13-15 of 63 FGQA – I8533 AD– 4.3.1 AQR to support FGQA available DV Testing Reports: PRC004097 Rev A, PRC001276 Rev A, PRC010923 Rev B, PRC002187 Rev A, PRC012445 Rev A, PRC045507Rev C | Specifications: 5 mm Trocar: DR000239 Rev J Sub System Requirement ID 0417 12 mm Trocar DR000249 Rev E System requirement ID SLR-07.11 & SLR-07.12 Reports: 5 mm Trocar: DV testing report PRC045507 Rev C 12 mm Trocar: QA0085 Rev L – Page 13-15 of 63 DV testing reports: PRC001276 Rev A PRC004097 Rev A | Specification: DR000164 Rev E-Requirement ID 2-67 for 5mm trocars & DR000164 Rev E-Requirement ID 2-69 for 11mm &12mm trocars Reports: Design Verification Report PRC012726 Rev B Page 17. Stack-up analyses Report PRC011851 Rev B. (Page 76,78 & 90) First Article and capability data met requirements: SCN003989 Rev B FGQA PR000471 Rev G- Sec 4.3.1 (Batch Identification) | Yes | As per Source specifications, the interface between the Obturator and the Sleeve is checked by Design Verification testing, under which the minimum clearance is calculated. The Finished Good Quality Assurance checks that the size mentioned on the product are appropriate and that the Obturator and Sleeve are of the same type and size, thus confirming that the devices would work as per Design Verification results, and additional simulation of actual operation is not necessary.
5.5.1 | Simulate the usage action and insert/extract the obturator in/from the puncture set for 3 times; the result shall comply with the provision of 4.5.1. | Visual Inspection of Size via FGQA and DV inspects the clearance at the interface of Obturator and Sleeve. | Visual Inspection of Size via FG Inspection report (PR001213 Rev B, PR001122 Rev C & PR001124 Rev F) and DV inspects the clearance at the interface of Obturator and Sleeve. | Visual Inspection of Size via FGQA. DV inspects the clearance at the interface of Obturator and Sleeve | No match but better and quantifiable test method | At source end, the requirement is verified using design verification tests and analysis based on the design specifications of the individual components (Obturator and Sleeve). These tests have quantifiable results, are stricter and more reliable in verifying the requirement and thus does not need additional repeated simulation of actual use.
4.5.2 | The maximum fit clearance (diameter difference) between puncture set cannula and obturator shall be no more than 0.3 mm. | Specification: DS0100-RevJ-Req. ID 008, DS0110-Rev K-Req. ID 008 DR000204-Rev F- Req. ID SYS-008 Reports: QA0081- Rev F – Page 11 of 44 QA0085 - Rev L- Page 13-15 of 63 DV Testing Reports: PRC004097 Rev A, PRC001276 Rev A, PRC010923 Rev B, PRC002187 Rev A, PRC012445 Rev A, PRC045507 Rev C | Specification: DR000239 Rev J Sub System Requirement ID 0417 DR000249 Rev E System requirement ID SLR-07.11 & SLR-07.12 Reports: DV testing report PRC045507 Rev C QA0085 REV L – Page 13-15 of 63 DV testing reports: PRC001276 Rev A PRC004097 Rev A | Specification: DR000164 Rev E-Requirement ID 2-67 for 5mm trocars & DR000164 Rev E-Requirement ID 2-69 for 11mm &12mm trocars Reports: Design Verification Report PRC012726 Rev B Page 17. Stack-up analyses Report PRC011851 Rev B. (Page 76,78 & 90) First Article and capability data met requirements: SCN003989 Rev B FGQA PR000471 Rev F- Sec 4.3.1 (Batch Identification) | Yes for 5 mm No for other sizes | Source specifications states the gap between obturator and the Sleeve to be 0.22m for 5mm trocars which is satisfying the requirement; however the gap for trocars of 8mm and above diameters is 0.533mm which is not meeting the requirement criteria of 0.3mm gap. Remarks: • The 0.5 mm clearance is clinically proven and vastly acceptable by physicians and there are no complaints of any kind related to the 0.5 gap in last 10 years. • A safety seal feature (on the top of the blade) ensures safety by retrieving blade once it penetrates abdominal wall and prevent tissue damage. This requires appropriate clearance to minimize drag for faster response of the safety feature. Otherwise with a small bending load, the safety cover can get stuck within sleeve and blade can remain open during insertion, which may damage the surgical area. • The 0.5 mm clearance provide a smooth insertion of obturator in the cannula and will not allow any tissue to get stuck during the procedure. • Also, the 0.5 clearance will provide sufficient space for fast flow of CO2 inside the abdominal cavity, when insufflation is required to be maintained while puncturing lever is still inside the cannula. • Unnecessary tight clearance requires tighter tolerance control on manufacturing of long parts, which may lead to unnecessary cost increase of components and hence device. | The maximum fit clearance (diameter difference) between puncture set cannula and obturator shall be no more than 0.3 mm 0.5 mm. Need actual measurements and Monte Carlo analysis, we have done. | Attachment 2 (Theoretical Analysis)
Attachment 10 (Test Data for TB12LT)
5.5.2 | Use a general/special measuring tool to make measurement. Take the maximum inner diameter of cannula and the minimum outer diameter of obturator according to the locations as shown in Figure 2; the diameter difference shall comply with the provisions of 4.5.2. | General or special measurement tool | General or special measurement tool | General or special measurement tool | Yes | Genreral Measurement tools are used to measure the dimensions.
4.5.3 | When the obturator is inserted into the puncture set completely, the tapered part of the head end of obturator shall be completely exposed. | Specification: DS0100-RevJ-Requirement ID 008, DS0110-Rev K-Requirement ID 008 DR000204-Rev F- Requirement ID SYS-008 Report: QA0081-Rev F – Page 11 of 44 QA0085 -Rev L– Page 13 of 63 EP3-0468 Rev A, EP-0469A, EP-0470A &EP-0483A 04-C-0016 PRC000843A PRC001082B PRC001281A | Specification: DR000239 Rev J Sub System Requirement ID 0417 DR000249 Rev E System requirement ID SLR-07.11 & SLR-07.12 Reports: DV testing report PRC045507 Rev C QA0085 Rev L – Page 13-15 of 63 DV testing reports: PRC001276 Rev A PRC004097 Rev A EP3-0468 Rev A, EP-0469 Rev A, EP-0470 Rev A &EP-0483 Rev A verified exposed cannula-length of no less than 0.0 mm (0.0 in) | Specification: DR000164 Rev E-Requirement ID 2-66 for 5mm trocars, DR000164 Rev E-Requirement ID 2-68 for 12mm &11mm trocars Report: Design Verification Report PRC012726 Rev B Page 16. Max.=0.1151 in (2.9mm). Min=0.0598 in (1.52mm) @ +/-5.1 Sigma (99.999966 % in tolerance) for 5 mm Max. =0.2370 in (6mm) Min. =0.1839 in (4.7mm) @ +/-5.1 Sigma for 11/12mm Trocar. Stack-up analyses Report PRC011851 Rev B. (Page 88 & 76) First Article and capability data met requirements: SCN003989 Rev B | Yes | Ethicon specifies this requirement as Obturator-Sleeve Transition: The interface of obturator’s cannula in the area Surrounding the sleeve-housings edge must be designed such that there is an exposed cannula-length of no less than 0.0 mm (0.0 in) parallel to the inside surface of the sleeve-housing (see figure below). Positive value of Dim X confirms that the Obturator bottom portion inlcuding the Obturator tip will always be extending outside the Sleeve. Functional Verification testings were conducted to verify this requirement. (Refer QA0085 - Pg 13-15 of 63) |  | Attachment 3
5.5.3 | Simulate the usage action, and make visual observation; the result shall comply with the provisions of 4.5.3. | General or special measurement tool | General or special measurement tool | General or special measurement tool | No match but better and quantifiable test method | Source specification uses special measurement tools to validate the requirement that the tip of Obturator should be exposed, when inserted in the Sleeve, at all times; whereas draft standard mentions "simulate the usage action and make visual observation". Source test is more reliable
4.6 | Connection Fastness
| If the sheath cap is detachable, it shall be firmly connected with the puncture set, while there shall be free from disconnection phenomenon on the joint. | Specification: DR000111 – Rev E- Req. ID 37, 116 DS0110- Rev K-Req. Id 032, 050 & DR000083- Rev H-Req. ID 025, 035 Reports: PRC001280A- 11mm and 12mm Outer Seal Related Design Verification Testing | Specification: 12 mm Trocar: DR000249 Rev E System Requirement ID SALR-07.3, SLR-08.5 Reports PRC053519 :Hermes EB3 - Force tojnstall Universal Cap, Actuate Latch Ring PRC058935 | There is no sheath cap in BASX Trocar, hence this clause is not applicable | Yes | Source specification defines the requirement as the force required to actuate the latching ring to attach or remove the Universal Cap from the sleeve housing shall be greater than 0.36 lbf and less than 6.84 lbf. with reliability of 99% at 95% Conf. interval. Integrated Removable Seal (Universal Reducer Cap) To Trocar Housing Attachment: The force to remove the Integrated Removable Seal (Universal Reducer Cap) from the sleeve housing of an ENDOPATH® Xcel 12 mm surgical trocar must be greater than 17.5 lbf. when loading is applied axial to the Trocar and the Integrated Removable Seal (Universal Reducer Cap). Rationale: The tensile axial force requirement is a function of the maximum trocar retention force for the Endopath XCEL Bladeless Trocar with stability sleeve. The requirement captures the required flexibility of assembly and disassembly of the sub-system, and quantifiable values which are set based on voice of customer, market surveillance and research makes the results stricter and more reliable in meeting the standards. Source requirement of minimum 17.5 lbf (78N approx.) is greater than the draft YY standard requirement of 50N, thus devices shall meet the requirement. |  | Attachment 4
5.6 | At the joint between the detachable sheath cap and the puncture set, apply axial static pulling force of 50 N on its any part for 10 s, and the result shall comply with the provisions of 4.6. | Test Method: Load Cells SCN019439B – Test method for force to install and removal of Integrated Removable Seal SCN019444 TR11513 | Test Method: Load Cells SCN019439B – Test method for force to install and removal of Integrated Removable Seal SCN019444 TR11513 | There is no sheath cap in BASX Trocar, hence this clause is not applicable | Yes | The cap is detached and attached with the use of latching mechanism. The force to assemble, disassemble and force of latching are all measured and tested and shall suffice the requirement. The Joint is also tested with axial load of 17.5lb(78N approx.) that meets the YY requirement (99.995% are greater than 17.712 lbf with 95% confidence)
4.7 | Sealing Performance and Air Impermeability
4.7.1 | The trocar shall have good sealing performance, with pressure drop not exceeding 0.1 kPa. | No corresponding requirement mentioned in the design requirement or specification documents. | No corresponding requirement mentioned in the design requirement or specification documents. | No corresponding requirement mentioned in the design requirement or specification documents. | No | 1. Pressure decline under 0.1kPa means that there will be no leak during the surgery once the initial pressure is achieved, indicating that there is no need of insufflator to supply the gas during the surgery. In actual working condition (surgery), an insufflator is provided to keep the desirable pressure in the abdominal area. These insufflators (globally) have insufflation rate between 8 to 14 L/min, which will help maintain insufflated abdominal volume to perform laparoscopic surgery without any fluctuation in the surgical space. 2. There is no need to keep the leak rate requirement so stringent because it will compromise other good features a surgeon is looking from the trocar while performing the surgery, such as: A zero-leak rate through 12 mm trocar seal can also be obtained through seal like deep cone seal used in 5 mm trocar but it won’t provide the other advantage a surgeon requires during the surgery such as: • Low drag force (insertion and withdrawal). If drag force is too high to exceed sleeve retention force, the trocar will be pulled out from patient and potentially cause patient injury and extend surgery time. • Use of a wide range of devices through the same seal. • Tissue retrieval with minimum leak and efforts • Better manipulation of device So, a trocar seal should be optimized while keeping all surgical needs (as listed below) in account • Maintain peritoneum cavity o Minimum Leak (accounting compensation by insufflator) • Less friction to devices o Less drag Force around seal (lubrication and less tight tolerances) • Allow a wide range of device sizes o System like Universal cap (seal can expand as per need of device size) • Allow tissue retrieval with minimum leaks o Size and Flexibility of sealing system (Universal cap) 3. The surgery room is usually air-conditioned to manage CO2 concentration and thus the CO2 leakage shall not become a concern. | The trocar shall have good sealing performance, with pressure drop not exceeding 0.1 kPa, leak rate not exceeding 0.5 L/min. | Attachment 9 (Pressure Drop Test on 5mm and 12mm Trocars)
4.7.2 | The trocar shall have good air impermeability, with pressure drop not exceeding 0.1 kPa. | No corresponding requirement mentioned in the design requirement or specification documents. | No corresponding requirement mentioned in the design requirement or specification documents. | No corresponding requirement mentioned in the design requirement or specification documents. | No |  | The trocar shall have good air impermeability, with pressure drop not exceeding 0.1 kPa, leak rate not exceeding 0.5 L/min. | Attachment 5 (Leak Rate test performed under DV testing)
5.7 | Conduct the test according to the method specified in Annex A; the sealing performance and air impermeability of the trocars shall comply with the provisions of 4.7. | Tests conducted as per draft YY standard at EES Labs. | Tests conducted as per draft YY standard at EES Labs. | Tests conducted as per draft YY standard at EES Labs. | Yes | Tests results suggest that although 5mm devices are comparatively easier to test, the 12mm devices show inconsistency in test results. The pressure drop limits are very stringent and does not appropriately capture the requirement. It shall be suggested to update the requirement to check leak rate of fluid from the seal in order to verify dealing performance and air impermeability | Refer tab "Annex A" of this document for suggested changes
4.8 | Fitting of gas injection valve
| The gas injection valve shall be equipped with Luer lock fittings, which shall comply with the provisions of GB/T 1962.2-2001. | Specifications: Stopcock valve Assembly Drawing: D13142 Rev K Valve Body- D12645 Rev L Raw Material Specification: R1002-13 G Cond A Valve Lever D12646 Rev N Raw Material Specification: R1017-36 Rev D Cond A Reports: CA055881 Rev A First Article Inspection report. Summary Report: SCN012930 Rev C “External Standards Review re: Luer connections in Endoscopy products”, | Specifications: Stopcock valve Assembly Drawing: D13142 Rev K Valve Body- D12645 Rev L Raw Material Specification: R1002-13 G Cond A Valve Lever D12646 Rev N Raw Material Specification: R1017-36 Rev D Cond A Reports: CA055881 Rev A First Article Inspection report. Summary Report: SCN012930 Rev C “External Standards Review re: Luer connections in Endoscopy products”, | Specifications: Stopcock valve Assembly Drawing: D13142 Rev K Valve Body- D12645 Rev L Raw Material Specification: R1002-13 G Cond A Valve Lever D12646 Rev N Raw Material Specification: R1017-36 Rev D Cond A Reports: CA055881 Rev A First Article Inspection report. Summary Report: SCN012930 Rev C “External Standards Review re: Luer connections in Endoscopy products”, | Yes | Conclusion : it is concluded that the original intent of the Luer fitting standard did not include applications such as trocars, and the standard was not intended to apply directly to trocars. Because of this, the full standard (in its entirety) does not apply to trocars and only portions of the standard will be applied as follows. (The original intent and scope of GB/T 1962.2-2001 standard was to specify requirements for luer lock fittings “for use with hypodermic syringes and needles and with certain other apparatus for medical use e.g. transfusion equipment.) Clauses 3.1, 3.2 & 4.1 of GB1962.2-2001 can be applied for the Trocar. Clause 3.1: Trocars utilize a female Luer lock connector, and thus the female fitting dimensions in Figure 1 apply (Refer Table 2 - SCN012930 Rev C) Clause 3.2: Trocars utilize female lock fittings that match the design shown in Figure 4 of the standard. (Refer Table 1- SCN012930 Rev C) Other clauses such as Clause 4.2-4.7 cannot be applied to Trocars because original intent of the standard did not include applications such as trocars, and the standard was not intended to apply directly to trocars Please refer SCN012930 Rev C page 4-7…for detailed analysis and reasoning on non-applicability of requirements of GB1962.2-2001 (EN1707) GB/T 1962.2-2001 refers ISO 594一2：1998 as its equivalent standard under the title heading on first page. The Forward section in ISO 594-2:1998 states that this standard corresponds to European Standard EN 1707:1996…. (reference SCN012930 Rev C (Notes Section… page 6 of 8) |  | SCN012930
5.8 | Conduct the test for Luer lock fitting of gas injection valve according to GB/T 1962.2-2001; the result shall comply with the provisions of 4.8. | Test Method: General or special measurement tool | Test Method: General or special measurement tool | Test Method: General or special measurement tool
4.9 | Penetration and insertion & pulling force
| The manufacturer shall perform evaluation on the penetration and insertion & pulling force of the trocars. | Specification: Penetration Force DS0110-009 DS0100-009 Peak Insertion force DS0100-021 DS0110-020 Sleeve Retention Force DS0100-017, 018 DS0110-016, 017 Withdrawal Drag Force DS0100-023 DS0110-022 Reports: EP3-0468 (B5LT) EP3-0469 (B5ST) PRC004097 / PRC003999 –(B5XT) PRC000843-B12LT PRC001281-B11LT PRC001082-B12LP PRC004097-B8LT PRC010923-B15LT EP3-0312 PRC000997 PRC056076,PRC056223 | Specification: Penetration Force DR000239 Rev J Customer Requirement ID 0359 DR000249 Rev E Performance Req. ID PLR-07.1 Peak Insertion Force DR000239 Rev J Customer Requirement ID 0318 DR000249 Rev E Performance Req. ID PLR-01.1 Retention Force DR000239 Rev J - SSID 0341 DR000249 Rev E - PLR-05.1 Withdrawal Force DR000239 Rev J - SID 0324 Reports for Penetration Force: Report - PRC045507 , 5mm Trocar PRC000843-B12LT PRC001281-B11LT PRC001082-B12LP PRC004097-B8LT PRC010923-B15LT Reports for Insertion Force: Report - PRC045507 , 5mm Trocar PRC000997 PRC056076,PRC056223 | Specifications: Penetration Force: DR000164 Rev E-Requirement ID 2-63, 64 Insertion DR000164 Rev E-Requirement ID 2-13, 16 Retention Force: DR000164 Rev E-Requirement ID 2-28, 29 Withdrawal Force: DR000164 Rev E-Requirement ID 2-31 Reports: Completion Report PRC012726 & Engineering Study PRC012954 Strategy Document PRC012725. | Yes | Ethicon specifies the penetration as a design requirement so as to have a safe, controlled, and predictable entry after a minimum of six penetration cycles, when evaluated in a porcine model . Trocar diameter < 5mm - FTP < 15lb Trocar diameter > 5mm - FTP < 25lb. Rational for 6 cycles was derived from study based on survey conducted on 14 surgens (Ref NTR-007) Also it specifies Peak Insertion Force considering 3 different scenarios as below: 1)The maximum allowable “Insertion” or forward drag force exerted on the 11mm clip applier, ER320, by the 11mm,12mm, or 15mm ENDOPATH® Xcel™ Bladeless Trocar system must not exceed 0.58 lbf. 2) In addition, the maximum allowable “Insertion” or forward drag force exerted on a 12.9 RMS 8 probe by the 12mm or 15mm ENDOPATH® Xcel™ Bladeless Trocar system must not exceed 1.49 lbf. 3) The maximum allowable “Insertion” or forward drag force exerted on a 11.6 RMS 8 probe by the 11mm ENDOPATH® Xcel™ Bladeless Trocar system must not exceed 1.02 lbf. Clarification: • As a rigorous standard, there should at least some bottom-line value to test with. • The force should be properly controlled: high insertion and withdrawal force will cause extra tearing to tissue and cause patient injury; also increase difficulty to use device and extend surgery time. | Change to 4.8. Puncture Force (or Force to penetrate FTP) The force that the surgeon must exert on the trocar to insert it into the surgical site should not exceed the following limits for a safe, controlled, and predictable entry 4.8.1 Trocar diameter 5 mm and less - FTP < 15lbf 4.8.2 Trocar diameter above 5 mm - FTP < 25lbf 4.9. Cannula (Sleeve) Retention Force (SRF) The force requires to remove the cannula from tissue should be greater that the following limits for stable holding of trocar/cannula during the procedure 4.9.1 Trocar diameter 5 mm and less – SRF > 7lbf 4.9.2 Trocar diameter above 5 mm - SRF > 14lbf 4.10 Device Insertion Force 4.10.1 Trocar diameter 5 mm and less < 0.31 lbf 4.10.2 Trocar diameter above 5 mm < 1.49 lbf 4.11 Device Withdrawl Force 4.11.1 Trocar diameter 5 mm and less < 0.33 lbf 4.11.2 Trocar diameter above 5 mm < 2.7 lbf | Attachment 6
Attachment 11 (12mm Trocar FTP and SRF measurement)
5.9 | The evaluation test methods for penetration and insertion & pulling force are given in Annex B and Annex C. Note: The methods provided in the annexes are only the test methods for evaluating the penetration and insertion & pulling force of the trocars. | Test Method PRC000843 : Wet Lab Procedure page 5-6 | Strategy - SCN007394 - Test Method - Drag Force | Strategy Document PRC012725. | Yes | Ethicon test method measures the penetration forces at Abdominal as well thoracic region in detailed way from reliability point of view. Results for the penetration force: B5LT: median = 13.245 lbf B12LT: Median insertion force into the thoracic cavity for the 12mm Endopath Xcel trocar (23.94 lbf) The median force to insert the Xcel 8mm trocar (10.045 lbf) The penetration force of the B15LT trocar (22.47 lbf) Results for Insertion Force 12.9 mm Probe – 93.5% of the population below 0.72 lbf with 95% confidence ER320 – 93.5% of the population below 0.34 with 95% confidence Peak insertion force 0.400 lbf for ER320 and 0.866 lbf for RMS 8 prob with 97.5% reliability @ 95% conf
4.1 | Corrosion Resistance
| For the components of trocar made of stainless steel, the corrosion resistance shall not be lower than the provisions of Class 6.6b in YY/T 0149-2006. | No corresponding requirement mentioned in the design requirement or specification documents since XCEL Trocars are developed for single patient use. | No corresponding requirement mentioned in the design requirement or specification documents since Optiview Trocars are developed for single patient use.. | No corresponding requirement mentioned in the design requirement or specification documentssince BASX Trocars are developed for single patient use.. | No | Disposable Abdominal Trocars are single patient use and are not designed for reuse. Thus, corrosion resistance requirement shall not be applicable for these devices. 1. The test specified in YY/T 0149-2206 are typically performed on devices and dental instruments that are re-sterilized and re-used. The Trocars are single-use devices and are not designed to be re-sterilized for multiple testing or use. Thus, as an example, performing a pressure steam test for corrosion is not an applicable test. 2. As part of the biocompatibility testing performed on these devices for ISO 10993, the devices are placed in saline for 72 hours at 50°C. The solution extract from the devices are then evaluated for biocompatibility and these devices passed this test. Additionally, the test lab evaluates the product and the solution for evidence of corrosion product or corrosion on the devices. In the tests on these devices, no observation was noted. This test at the elevated temperature is an extreme condition for a single-use device and would indicate that no corrosion would occur in the intended use of these devices. 3. Post Market Surveillance reports did not report any complaints for these devices related to corrosion in the past seven (7) years | The requirement shall be suggested to be removed for disposable trocars.
5.1 | Conduct the test for corrosion resistance according to the provisions specified in 6 sodium chloride solution test method of YY/T 0149-2006; the result shall comply with the requirements of 4.10. | No corresponding requirement mentioned in the design requirement or specification documents. | No corresponding requirement mentioned in the design requirement or specification documents. | No corresponding requirement mentioned in the design requirement or specification documents. | Yes | Tests are currently being performed in EES Labs as per the test method mentioned in draft YY standard to verify this requirement.
4.11 | Sterility
| The trocar shall be sterilized though a validated sterilization process, while the sterilized trocar shall be sterile. | Specification: DS0110-Rev K- Req. ID 071-074 DS0100-Rev J- Req. ID 070-074 Report: Gamma Sterilized (Cobalt 60) SCN001051 RevL:Sterilization Strategy SCN019558 Rev X: Sterilization Strategy SCN000338 Rev Z : Sterilization Strategy | Specifications: DR000239 Rev J Customer Requirement ID 0485 DR000249 Rev E Customer Requirement ID CLR-15 Reports 5 mm Trocars: Gamma Sterilized (Cobalt 60) Sterilization Report SCN012729 Rev H: Gamma Sterilized (Cobalt 60) Sterilization Report SCN019558 Rev X: | Specification: DR000164 Rev E- Customer Req Id- 2-97 Report: PRC013145 Rev A-B&F Test Protocol PRC013146 Rev A-B&F Test Report Gamma Sterilized (Cobalt 60) SCN003244 Rev J: Sterilization Strategy | Yes | The devices are sterilized by Gamma radiation (Cobalt 60) as per ISO11137-2:2006 which is equivalent to US and Chinese Pharmacopoeia
5.11 | Conduct the test according to the "sterility test method" specified in Chinese Pharmacopoeia (Edition 2015, Volume IV); the test result shall comply with the provisions of 4.11. | Sterilization Testing as per US Pharmacopoeia. ANSI/AAMI/ISO11137:2006 ANSI/AAMI/ISO11737-1:2006 ANSI/AAMI/ISO11737-2:2009 | Sterilization Testing as per US Pharmacopoeia. ANSI/AAMI/ISO11137:2006 ANSI/AAMI/ISO11737-1:2006 ANSI/AAMI/ISO11737-2:2009 | Sterilization Testing as per US Pharmacopoeia. ANSI/AAMI/ISO11137:2006 ANSI/AAMI/ISO11737-1:2006 ANSI/AAMI/ISO11737-2:2009 | Yes | The devices are sterilized by Gamma radiation (Cobalt 60) as per ISO11137-2:2006 which is equivalent to US and Chinese Pharmacopoeia
4.12 | Ethylene Oxide Residue
| If the trocar is sterilized by ethylene oxide, the ethylene oxide residues shall not be greater than 10 μg/g. | PRC002494 Rev A Appendix S (Page 55 of 56) Min EO (μg/g) = 27.8 | Specification: DR000239 Rev J Requirement ID -0486 Reports: PRC002494 Rev A Appendix S (Page 55 of 56) Min EO (μg/g) = 27.8 (Equivalency claimed for all material other than the adsorbent being tested by SCN007937) | Specification: DR000164 Rev E- Customer Req Id- 2-101 Reports: PRC002494 Rev A Appendix S (Page 55 of 56) Min EO (μg/g) = 27.8 | No | EO requirement is not met by XCEL Trocars | Change “the ethylene oxide residue shall not be bigger than 10 μg/g” to “the ethylene oxide residue shall meet ISO10993 (GB16886)” | Attachment 7 (Available Test Method at EES for EO Residue Testing)
5.12 | Conduct the test according to the gas chromatographic method specified in GB/T 14233.1-2008; the result shall comply with the requirements of 4.12.
4.13 | Packaging, identification and instructions for use
4.13.1 | The single package of trocar shall bear the identification that complies with the provisions of 3.1.2. | Could not find nomenclature as per 3.1.2 | Could not find nomenclature as per 3.1.2 | Could not find nomenclature as per 3.1.2 | No | The indication of Trocar should be a representative dimension for device sizing not the nominal Trocar ID. Surgeons use this indication on label to select a device size, which can be inserted in the trocar. Precise nominal ID requirement on label or IFU may lead to user confusion and unnecessary label changes. | The requirement shall only have the mention for the device specification code which helps to match the suitable device with the size of the trocars and the working/effective length of the Trocar. | Attachment 16 (Reference Design Spec Artwork on Tyvek)
4.13.2 | The instructions for use of trocar shall contain the contents of the specifications for the devices that may be used in combination with the trocar. | Specification: DS0110-Requirement ID 054-062 Evidence: A88361-IFU Bladeless Trocar A87204-IFU Blunt Tip Trocar A88363-IFU Bladed Trocar A88389- IFU Universal Sleeve | Instruction for Use A87795, A87796 A87797 | Instruction for Use A88361-IFU BASX Trocar | Yes | Instruction for use covers YY requirement
5.13 | Conduct visual inspection and the result shall comply with the provisions of 4.13. | Test Method : Visual | Test Method : Visual | Test Method : Visual | Yes | Visual inspection meets the requirement as per YY
4.14 | Biocompatibility
| The biological evaluation shall be performed for the trocar according to the requirements of GB/T 16886.1; the result shall indicate that there is no biological hazard. | Reports: DBAM: SCN001051 | Reports: DBAM SCN019962 Rev E: - DBAM SCN027042 Rev C | Reprots DBAM SCN003244 Rev J | Yes | Biological evaluation is performed as per ISO 10993-1 standard for medical Devices and The same ISO 10993-1 is referenced/used by China standard.
5.14 | The biological evaluation shall be performed for the trocar according to the requirements of GB/T 16886.1; the result shall indicate that there is no biological hazard. | Reports: DBAM: SCN001051 | Reports: DBAM SCN019962 Rev E: - DBAM SCN027042 Rev C | Reprots DBAM SCN003244 Rev J | Yes | Biological evaluation is performed as per ISO 10993-1 standard for medical Devices and The same ISO 10993-1 is referenced/used by China standard.

## 7 Packaging - 8 Identification
Clause No. | Draft YY Requirement/Testing | Supporting Evidence
7 | Packaging
7.1 | The manufacturer shall be able to provide the provisions to prove that the package loaded with trocar complies with the requirements of GB/T 19633.1. | CP0154 refers to ISO 11607 which is equivalent to GB/T 19633.1
7.2 | If it is sterilized by ethylene oxide, dialysis packaging shall be used for packaging. | The device is not sterilized by ethylene oxide. This clause is not applicable
7.3 | The marks of opening shall be remained on package once opened. | WE0154 - section 4.4.2 - Tamper evidence
8 | Identification for single package
At least the following information that can be clearly identified shall be provided on single package.
| a) Product name; | Specification; All package assembly drawings (eg. P71810) mentioned in page 46 of 74 of Technical file TF-N000010 All Label carton artworks (eg. eg. A95484) from the package assembly drawings. MS00003: 4.5.1 Package should have correct printing color. 4.5.2 Package should not have stray ink. Page 28 of 48 - Labelling information 4.10.1 All print located on the primary label shall be present. 4.10.2 Product code, description, or graphic of the device on primary label shall be correct, present, legible, and undamaged. 4.10.3 Shelf life printed on primary label shall not exceed validated shelf life. 4.10.4 Shelf life printed on primary label and shall not be less than validated shelf life. 4.10.5 Shelf life printed on primary label shall be present, legible, and undamaged. 4.10.6 Device size on primary label shall be correct, present, legible, and undamaged. 4.10.7 Package batch information on primary label shall be correct, present, legible, and undamaged. 4.10.8 Logo (marketed name of device or trademark) on primary label shall be correct, present, legible, and undamaged. 4.10.9 Manufacturer’s name and contact information on primary label shall be correct, present, legible, and undamaged. 4.10.10 Barcode on primary label shall be correct, present and scanable. 4.10.12 Point of Use Labels should not have marred, smeared, or missing print on any portion of the label. 4.17.1 Safety related markings and symbols (see Appendix I) shall be correct. 4.17.2 Safety related markings and symbols (see Appendix l) shall be present, legible, and undamaged. 4.17.3 Non-safety related markings and symbols (see Appendix I) shall be correct. 4.17.4 Non-safety related markings and symbols (see Appendix l) shall be present, legible, and undamaged. Report: Packaging line FGQA : Xcel Trocar I8533 Rev AD, PR000316- Rev AB section 4.3 Optiview 5 mm Trocar PR001213 Rev B Section 4.1 12 mm Trocar PR001122 Rev C Section 4.1 12 mm Trocar (2D12LT/2D12XT) PR001124 Rev F Section 4.1 Package stability test - SCN003465 PRC045504, and PRC045505 are Real-time Aging Study and Completion Report
b) Identification for product specification;
c) Name and address of manufacturer and/or distributor;
d) Production date and/or batch number;
e) Shelf life or expiration date;
f) Sterilization method;
g) It is provided in sterile state and for single use;
h) Instructions for use and precautions, including the warning of “Do not use when package is broken”;
i) Indicate the quantity for non-single package;
Note: The graphic symbols provided in YY/T 0466.1 can be used to comply with the corresponding requirements mentioned above.

## Annex A
Annex A
Test method for sealing performance and air impermeability
(Normative) | Suggestions: | Justification/Remarks:
| Change the test method to following: 5.7.1 Select the device with the minimum specification as specified in the instructions for use or the corresponding standard to insert into the puncture set, close the gas injection valve, and introduce air pressure (higher than atmospheric pressure by 2 kPa) from the head end of puncture set for continuous 30 s. The result shall comply with the provisions of 4.7.1. 5.7.2 After the test specified in 5.7.1, extract the puncturing lever, introduce air pressure (higher than atmospheric pressure by 2 kPa) from the head end of puncture set for continuous 30 s. The result shall comply with the provisions of 4.7.2. | 1. 4kPa is unreasonable because insufflator pressure is set within 9-18 mmHg (1.2-2.4kPa) according to global market research, including US, EU and Asia. 2. This method (as mentioned in Annex A) is of low stability because it is vulnerable to environment turbulence.
A.1 Definition
Sealing performance and air impermeability: Evaluate the pressure loss when the obturator of trocar and the standard rod of simulating endoscopic device are inserted into/pulled out from the puncture cannula.
A.2 Test apparatus
A.2.1 Pressure device, including air source, pressure gauge (with the precision of 0.1 kPa) and connecting pipe.
(Note: The volume of the entire pressure device system does not exceed 30 mL)
A.2.2 The requirements for standard large/small rod are shown as follows:
a) Solid stainless steel rod
b) The diameter shall be equivalent to the diameter of the device rods of the minimum and the maximum specifications that are applicable to trocars, with the error of ±0.1 mm
c) The surface roughness Ra shall not be greater than 0.4 μm
d) The puncture end shall be in circular end shape
e) The length shall not be less than the working length of the obturator
A.3 Test methods
A.3.1 Connect the head of the trocar to the pressure device, so as to ensure that there is no leakage at the joint.
A.3.2 Close the gas injection valve of the trocar, and block the side hole of the puncture cannula and the air hole at the tail end of the obturator.
A.3.3 Apply gas pressure of 4 kPa to the head of the trocar by means of a pressure device, and stop pressurization after the pressure has become stable for 10 s.
A.3.4 1 min later, record the changes in pressure gauge.
A.3.5 Pull out the obturator, apply pressure to stabilize the pressure at 4 kPa, and stop pressurization after the pressure has become stable for 10 s
A.3.6 1 min later, record the changes in pressure gauge.
A.3.7 Replace the obturator with a standard large/small rod, repeat the steps of A.3.1~A.3.6, and record the changes in the pressure gauge, respectively.

## Annex B
Annex B
Test method for evaluation of penetration force
(Recommendatory)
B.1 Definition
Penetration force: The maximum force required for the trocar to penetrate the test material at a certain speed.
B.2 Test equipment and test material
B.2.1 Material testing machine, with the accuracy of 0.1 N
Parameter setting: puncture speed of 100 mm/min
B.2.2 Fixture is used to ensure that the trocar and the test material are firmly fixed. The lower part of the puncture site is supported by a circular block. When sufficient support is provided for the object to be punctured, it is ensured by the circular middle space that the puncture core can completely penetrate the object to be punctured.
B.2.3 Test material: abdominal tissue of pig, with skin, preserved in chilled condition for no more than 48 h. Unilateral size is at least 10 times the outer diameter of the puncture cannula.
B.3 Procedures
B.3.1 Correctly install and clamp the trocar and the test material (B.2.3).
B.3.2 On the epidermis of test material (B.2.3), mark out a cut with the size being equivalent to the diameter of the cannula, and the dermis layer shall be completely cut open.
B.3.3 Adjust the position of the upper fixture and confirm that the head of the trocar is in contact with the cut of test material (B.2.3) but is not under stress.
Note: The direction of cutting edge on the obturator shall be consistent with that of cut on the test material.
Start, use the trocar to puncture the test material (B.2.3), until the thread structure on the puncture cannula completely enters the dermis layer of test material (B.2.3), and then stop. The recorded maximum force is determined as the penetration force.

## Annex C
Annex C
Test method for evaluation of insertion & pulling force
(Recommendatory)
C.1 Definition
Insertion/extraction force: The maximum force required for the pressure loss when the obturator of trocar and standard large/small rod are inserted into/pulled out from the puncture cannula.
C.2 Test equipment
C.2.1 Material testing machine, with the accuracy of 0.1 N
Parameter settings: puncture/pull-out speed of 100 mm/min, puncture stroke of L-10 mm (if the trocar structure cannot meet L-10 mm, it is allowed to modify the stroke, which shall be indicated in the test report)
C.2.2 It can be ensured by the fixture that the puncture cannula is not subjected to stress in the radial direction and is firmly fixed in the axial direction.
C.3 Procedures
C.3.1 Correctly install and clamp the puncture cannula and the obturator to ensure that the axis of the obturator coincides with the axis of the puncture cannula.
C.3.2 Adjust the position of the upper fixture to confirm that the head of the obturator is in contact with the sealing ring on the puncture cannula but is not subjected to stress, as shown in Figure C.1.
Note: If there is a positional correspondence between the obturator and the cannula, it is required to ensure that the relative position between the puncture core and the cannula is correct.
C.3.3 Start, insert the obturator into the puncture cannula until it reaches the preset stroke position, and then stop; the recorded maximum force value is determined as the insertion force of the obturator, as shown in Figure C.2.
C.3.4 At the position where the test in C.3.3 is completed, continue to pull out the obturator at the preset speed, until it is separated from the puncture cannula, while the recorded maximum force value is determined as the extraction force of obturator.
Figure C.1 Figure C.2
C.3.5 Use the obturator to replace the standard large/small rod (A.2.2), repeat the steps of C.3.1~C.3.3, and the recorded maximum force value is determined as the insertion force of the standard large/small rod; repeat the step of C.3.4, and the recorded maximum force value is determined as the extraction force of the standard large/small rod.