Authored By:  Mark Glassett

TABLE OF CONTENTS	Page

1.	ABSTRACT	1

2.	OBJECTIVE	1

3.	RESULTS	2

4.	DISCUSSION	2

5.	CONCLUSIONS	3

6.	RECOMMENDATIONS	3

# ABSTRACT

Zip Pen samples that have been exposed to three year accelerated aging, shipping cycle temperature extremes and the ship test were subjected to the dielectric withstand tests required for compliance to IEC 60601-2-2: 2009. The Zip Pen passed the requirements of the standard for these tests. The specific clause of the IEC 60601-2-2 standard for each test is listed below. In the original testing (report ENG-RPT-328), there were Mains Dielectric test failures. This report documents the retest of the product after design improvements and is for the HF and Mains dielectric only. This testing was performed per protocol ENG-PRT-283. The original protocol was ENG-PRT-227.

| Required Test | IEC 60601-2-2 Clause | Report Reference |
| --- | --- | --- |
| Zip Pen HF Dielectric | 201.8.8.3.103 | This Report |
| Zip Pen Mains Dielectric | 201.8.8.3.104 | This Report |
| Cable Leakage Current | 201.8.8.3.102 | ENG-RPT-328 |
| Cable HF Dielectric | 201.8.8.3.103 | ENG-RPT-328 |
| Cable Mains Dielectric | 201.8.8.3.104 | ENG-RPT-328 |
| Plug Mains dielectric | 201.8.8.3.104 | ENG-RPT-328 |

# OBJECTIVE

The objective of this test report is to document compliance of the Zip Pen catalog items 2525-10 and 2525-15 with IEC60601-2-2: 2009 requirements for Dielectric Withstand after accelerated aging to simulate three years expiration life, shipping extremes and ship testing. Note that for the electrical tests, the 2525-10 and 2525-15 are considered equivalent. The only difference between these two products is cable and tubing length. None of the dielectric withstand testing is affected by the cable or tubing length.

# RESULTS

## Zip Pen Handpiece High Frequency Dielectric Withstand

Thirty samples of the Zip Pen were tested for high frequency dielectric withstand. The test voltage was 6,600 Vp which is the rated accessory voltage of 5,500 Vp plus 20% as required by the standard. All samples passed the HF dielectric test requirements of the protocol. This data is shown in Appendix I.

## Zip Pen Handpiece Mains Frequency Dielectric Withstand

The Mains frequency dielectric test followed the HF dielectric test on the same samples as required by the standard. The test voltage was 4,600 Vrms. This is derived from the standard that requires the test voltage to be the rated peak voltage plus 1,000 peak volts and converted to Vrms. All samples passed the Mains frequency dielectric test requirements of the protocol. This data is shown in Appendix I.

## Zip Pen Handpiece Continuity Measurement

Thirty samples of Zip Pen were tested for button continuity less than 50 ohms when the button is depressed and greater than 10,000 ohms when the button is not depressed. Both Cut and Coag buttons were tested and passed the requirements of the protocol. Test data is shown in Appendix I.

# DISCUSSION

The samples used for the testing were catalog 2525-15 Zip Pen Lot #S140286 sterilized to a minimum gamma exposure of 50 kGy. See appendix II for the product C of C and sterilization record. The 2525-10 and 2525-15 catalog numbers are the same product with the exception of tube and cord length. As previously mentioned, the Zip Pen had previously failed Mains frequency dielectric and this is a retest after a design change. The design change consisted of the addition of another layer of tape insulation over the PCB and Dome Switch area. The additional layer of tape is a Surlyn material chosen for its dielectric properties and gamma sterilization stability. The addition of this tape to the assembly is shown on drawing ENG-DWG-716 rev 002 and ENG-DWG-717 rev 002.

Prior to testing, these samples were subjected to accelerated aging at 55˚C to simulate three years shelf life, see aging process sheet in Appendix II. They were also exposed to temperature and humidity extremes per protocol ENG-PRT-283 to simulate possible exposure during the shipping and storage of the device. For documentation of this conditioning see shipping test log sheet in appendix II (note that the shipping test log sheet is from protocol X1150720-10). The sample size required for each test was 30 units.

## Zip Pen Handpiece High Frequency Dielectric Withstand

The Zip Pen has a voltage rating of less than or equal to 5,500 Vp. The High Frequency dielectric test was performed at minimum of 6,600 Vp which is 120% of the rated voltage as required by IEC 60601-2-2 clause 201.8.8.3.103. The testing followed the requirements of protocol ENG-PRT-283 and all of the 30 samples passed the requirements of the protocol.

## Zip Pen Handpiece Mains Frequency Dielectric Withstand

## The Mains Frequency dielectric test was performed at 4,600 Vrms which is 1000 Vp above the rated voltage as required by IEC 60601-2-2 clause 201.8.8.3.104. This value is calculated by adding 1,000 volts to the rated accessory voltage of 5,500 Vp and dividing by √2 to convert the voltage to Vrms. The testing followed the requirements of protocol ENG-PRT-283 and all of the 30 samples passed the requirements of the protocol.

## Zip Pen Handpiece Continuity

The Zip Pen continuity was tested both prior to and after the dielectric test. The continuity of all samples was less than 50 ohms when each button was depressed and read open circuit when no buttons were depressed.

# CONCLUSIONS

This testing in addition to the testing reported in ENG-RPT-328 demonstrates that the Zip Pen catalog numbers 2525-10 and 2525-15 comply with the required standards referenced in section 1 after three year accelerated aging and shipping and storage conditions.

# RECOMMENDATIONS

This testing was performed to demonstrate compliance of the Zip Pen to IEC 60601-2-2: 2009 after three year accelerated aging to support the three year expiration life. This accelerated age test will support the three year expiration life of the product for market introduction. Real time age samples from the first production lot will be put aside for testing per Megadyne Protocol ENG-PRT-057.

The testing also demonstrated compliance to IEC 6060 1-2-2: 2009 after exposure to extreme shipping and storage conditions. The shipping box labels of the products will show the international symbols for shipping and storage with temperatures of 5°C to 50°C and relative humidity of 15% to 95%. The IFU will include the note “Normal storage conditions are assumed. Brief excursions to limits allowed”.

APPENDIX I

Zip Pen Handpiece High Frequency and Main Frequency Dielectric

APPENDIX II

Zip Pen Aging, C of C, Sterilization, Ship Test Record

APPENDIX II

Zip Pen Aging, C of C, Sterilization, Ship Test Record

APPENDIX II

Zip Pen Aging, C of C, Sterilization, Ship Test Record

APPENDIX II

Zip Pen Aging, C of C, Sterilization, Ship Test Record

APPENDIX II

Zip Pen Aging, C of C, Sterilization, Ship Test Record

APPENDIX II

Zip Pen Aging, C of C, Sterilization, Ship Test Record